## Blank Template
A L T A M E D I C A I N C O R P O R A T E D Medical Products You Can Trust!
CERTIFICATE OF CALIBRATION
Certificate Number: |  | ALMC-ZIPIQ-TT-12909-022024 | [5 empty] | 2024-02-25
Performed by:
Company Name: |  | Altamedica Incorporated
Address: |  | 199R Dicen St., Brgy. Bantud, Lapaz, Iloilo City
Contact Number: |  | Tel#: (033) 508 0208 / (033) 333 1769, Fax#: (033) 508 0208
Customer:
Company Name:
Address: |  | Cebu City
Contact Number: |  | .
Equipment:
Equipment Type: |  | Clinical Centrifuge - 6 placer |  |  | Manufacturer: |  | LW Scientific
Model: |  | ZIPIQ-TT |  |  | Manufacture Date: |  | .
Serial Number: |  | 6M18112909 |  |  | Distributor: |  | Altamedica Incorporated
As received: |  | In good condition. |  |  | As Released: |  | In good condition.
Test Condition:
Room Temperature: |  | 25°C |  |  | Relative Humidity: |  | 40%
Performed at: |  | Altamedica Incorporated
Methods and Materials:
Method(s): |  | Non-contact speed measurement
Procedure(s): |  | Centrifuge Work Instruction Manual
Material(s) used: |  | YX360TRF Sanwa Multimeter |  |  | Serial Number: |  | N/A
|  | TKRT 20 SKF Tachometer |  |  | Serial Number: |  | 1818016
Calibration Result:
Calibration Date: |  | 2024-02-25 |  |  | Calibration Due: |  | 2025-02-25
Speed Test: |  | Passed |  |  | Timer Test: |  | Passed
Conditions/Analysis: |  | Machine is in good working condition. All results are within acceptable range.
This report may not be reproduced, except in full, without the written approval of the Technical Manager. The results reported herein apply only to the calibration of the item described above, and no limitations of use apply to the calibrated unit. Although the calibrated item meets the specifications and performance at the time of calibration, due to any number of factors, the recommended due date of the item calibrated does not imply continuing confidence to specifications during the recommended interval.
Performed by: | [6 empty] | Noted by:
Joanna Faye B. Talagon | [6 empty] | Engr. Rey Christopher Alipe
Service Engineer | [6 empty] | Technical Manager
ANALYTICAL DATA
Certificate Number: ALMC-ZIPIQ-TT-12909-022024
2024-02-25
Speed Test
| Speed (rpm) |  |  |  |  | Allowed Tolerance |  | Remarks
| Set Value |  | Displayed Value | Measured Value
| 5000 rpm |  | 1250 | n/a |  | 125 |  | n/a
|  |  | 2500 | n/a |  | 250 |  | n/a
|  |  | 3750 | n/a |  | 375 |  | n/a
|  |  | 5000 | 5008 |  | 500 |  | Pass
Timer Test
| Time (sec) |  |  | Difference |  | Allowed Tolerance |  | Remarks
| Set Time |  | Measured Time
| 10 |  | 10 | 0 |  | 1 |  | Pass
| 60 |  | 60 | 0 |  | 6 |  | Pass
| 120 |  | 120 | 0 |  | 12 |  | Pass
| 300 |  | 300 | 0 |  | 30 |  | Pass
Remarks
Pass! Machine is ready to use.
Performed by: | [6 empty] | Noted by:
Joanna Faye B. Talagon | [6 empty] | Engr. Cris Ian H. Montero
Service Engineer | [6 empty] | Technical Manager
Note on C5: CERT NUMBER
Note on J5: Date Here!
Note on C29: Pass or Fail
Note on I29: Pass or Fail
Note on C30: REMARKS!