## Blank Template
A L T A M E D I C A I N C O R P O R A T E D Medical Products You Can Trust!
CERTIFICATE OF CALIBRATION
Certificate Number: |  | -DSB500-waterbath-0012-062025 | [5 empty] | June 4, 2025
Performed by:
Company Name: |  | Altamedica Incorporated
Address: |  | 199R Dicen St., Brgy. Bantud, Lapaz, Iloilo City
Contact Number: |  | Tel#: (033) 508 0208 / (033) 333 1769, Fax#: (033) 508 0208
Customer:
Company Name: |  | .
Address: |  | .
Contact Number: |  | .
Equipment:
Equipment: |  | Water Bath |  |  | Manufacturer: |  | Digisystem Laboratory Instruments
Model: |  | DSB-500 |  |  | Manufacture Date: |  | N/A
Serial Number: |  | SB1-25050012 |  |  | Distributor: |  | Altamedica Incoporated
As received: |  | In good condition |  |  | As Released: |  | In good condition
Test Condition:
Room Temperature: |  | 25 |  |  | Relative Humidity: |  | 45%
Performed at: |  | Altamedica Incorporated
Methods and Materials:
Method(s): |  | Water Blank Method
Procedure(s): |  | Semi Automated Chemistry Analyzer Work Instruction Manual
Material(s) used: |  | Description |  |  | Serial Number / Lot Number |  |  | Expiry
|  | Fluke 51-2 Thermometer |  |  | 36370220WS |  |  | N/A
Calibration Result:
Calibration Date: |  | June 3, 2025 |  |  | Calibration Due: |  | June 3, 2025
Stability: |  | Pass |  |  | Linearity: |  | Pass
Conditions/Analysis: |  | Machine is in good condition
This report may not be reproduced, except in full, without the written approval of the Technical Manager. The results reported herein apply only to the calibration of the item described above, and no limitations of use apply to the calibrated unit. Although the calibrated item meets the specifications and performance at the time of calibration, due to any number of factors, the recommended due date of the item calibrated does not imply continuing confidence to specifications during the recommended interval.
Performed by: | [6 empty] | Noted by:
Joanna Faye B. Talagon | [6 empty] | Engr. Rey Christopher A. Alipe
Service Engineer | [6 empty] | Technical Manager
Calibration Data
Certificate Number: -DSB500-waterbath-0012-062025
June 4, 2025
Pre - Calibration
I. Stability
Target Temperature: |  | 37 |  |  | Tolerance: |  | 0.5
Repeat No. |  | Time |  | Measured Temperature |  | Error |  | Remarks
1 |  | 4:10 |  | 37.3 |  | 0.3 |  | Pass
2 |  | 4:15 |  | 37 |  | 0 |  | Pass
3 |  | 4:20 |  | 36.9 |  | 0.1 |  | Pass
4 |  | 4:25 |  | 36.9 |  | 0.1 |  | Pass
5 |  | 4:30 |  | 37 |  | 0 |  | Pass
II. Linearity
Maximum Working Temperature: |  |  | 37 |  | Tolerance: |  | 0.2
No. | Target Temperature |  | Measured Temperature |  | Error |  | Remarks
1 | 4 |  | 3.9 |  | 0.1 |  | Pass
2 | 12 |  | 12 |  | 0 |  | Pass
3 | 24 |  | 24.1 |  | 0.1 |  | Pass
4 | 36 |  | 36 |  | 0 |  | Pass
5 | 48 |  | 48 |  | 0 |  | Pass
Post - Calibration
I. Stability
Target Temperature: |  | 37 |  |  | Tolerance: |  | 0.5
Repeat No. |  | Time |  | Measured Temperature |  | Error |  | Remarks
1 |  | 17:08 |  | 37.3 |  | 0.3 |  | Pass
2 |  | 17:13 |  | 37 |  | 0 |  | Pass
3 |  | 17:18 |  | 36.9 |  | 0.1 |  | Pass
4 |  | 17:23 |  | 36.9 |  | 0.1 |  | Pass
5 |  | 17:28 |  | 37 |  | 0 |  | Pass
II. Linearity
Maximum Working Temperature: |  |  | 37 |  | Tolerance: |  | 0.2
No. | Target Temperature |  | Measured Temperature |  | Error |  | Remarks
1 | 4 |  | 4 |  | 0 |  | Pass
2 | 12 |  | 12 |  | 0 |  | Pass
3 | 24 |  | 24 |  | 0 |  | Pass
4 | 36 |  | 36 |  | 0 |  | Pass
5 | 48 |  | 48 |  | 0 |  | Pass
Remarks
Machine has been calibrated on its operating temperature. Machine is ready to use.
Performed by: | [6 empty] | Noted by:
Joanna Faye B. Talagon | [6 empty] | Engr. Rey Christopher A. Alipe
Service Engineer | [6 empty] | Technical Manager
Note on C5: CERT NUMBER
Note on J5: Date Here!
Note on C32: Pass or Fail
Note on I32: Pass or Fail
Note on C33: REMARKS!
Note on C52: Military Time: EX. 13:30
Note on C70: Military Time: EX. 13:30