## Blank Template
A L T A M E D I C A I N C O R P O R A T E D Medical Products You Can Trust!
CERTIFICATE OF CALIBRATION
Certificate Number: | [7 empty] | Date:
Performed by:
Company Name: |  | Altamedica Incorporated
Address: |  | 2nd Floor Masonic Temple Bldg. J.M. Basa Street, Iloilo City
Contact Number: |  | Tel#: (033) 508 0208 / (033) 333 1769, Fax#: (033) 508 0208
Customer:
Company Name:
Address:
Contact Number:
Equipment:
Equipment: | [5 empty] | Manufacturer:
Model: | [5 empty] | Manufacture Date:
Serial Number: | [5 empty] | Distributor:
As received: | [5 empty] | As Released:
Test Condition:
Room Temperature: | [5 empty] | Relative Humidity:
Performed at:
Methods and Materials:
Method(s): |  | Tri-Level Control Calibration
Procedure(s): |  | Hematology Analyzer Work Instruction Manual
Material(s) used: |  | YX360TRF Sanwa Multimeter |  |  |  | Serial Number:
|  | Tri-Level Hematology Control |  |  |  | Brand:
Lot Number:
Expiration Date:
Calibration Result:
Calibration Date: | [5 empty] | Calibration Due:
Background Test: |  |  |  | Repeatability: |  |  |  | Accuracy:
Conditions/Analysis:
This report may not be reproduced, except in full, without the written approval of the Technical Manager. The results reported herein apply only to the calibration of the item described above, and no limitations of use apply to the calibrated unit. Although the calibrated item meets the specifications and performance at the time of calibration, due to any number of factors, the recommended due date of the item calibrated does not imply continuing confidence to specifications during the recommended interval.
Performed by: | [7 empty] | Noted by:
Engr. Cris Ian H. Montero
Service Engineer | [7 empty] | Technical Manager
ANALYTICAL DATA
Certificate Number:
Background Test
|  | Parameter |  | WBC | RBC | HGB | HCT | PLT
|  | Unit |  | x10⁹/L | x10¹²/L | g/L | % | x10⁹/L
|  | Acceptable Range |  | 0.2 | 0.02 | 1 | 0.5 | 10
Background Test
|  | Remarks |  | Pass | Pass | Pass | Pass | Pass
Repeatability
|  | Parameter |  | WBC | RBC | HGB | MCV | PLT
1
2
3
4
5
|  | MEAN |  | #NAME? | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | SD |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | CV (%) |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | Allowed CV (%) |  | 2 | 1.5 | 1.5 | 0.4 | 4
|  | Remarks |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Pre - Calibration
I. Coefficient
|  | Parameter |  | WBC | RBC | HGB | MCV | PLT
Coefficient (%)
|  | Coeff. Range (%) |  | 75 - 125 | 76 - 125 | 77 - 125 | 78 - 125 | 79 - 125
|  | Remarks |  | Fail | Fail | Fail | Fail | Fail
II. Control Data
Lot Number: | [6 empty] | Expiration Date:
| Level | Unit | Para. | Measured | Target | Range | Error | Remarks
| Low | x10⁹/L | WBC |  |  |  | 0 | No Data
|  | x10¹²/L | RBC |  |  |  | 0 | No Data
|  | g/L | HGB |  |  |  | 0 | No Data
|  | x10⁹/L | PLT |  |  |  | 0 | No Data
|  | % | HCT |  |  |  | 0 | No Data
|  | fL | MCV |  |  |  | 0 | No Data
| Normal | x10⁹/L | WBC |  |  |  | 0 | No Data
|  | x10¹²/L | RBC |  |  |  | 0 | No Data
|  | g/L | HGB |  |  |  | 0 | No Data
|  | x10⁹/L | PLT |  |  |  | 0 | No Data
|  | % | HCT |  |  |  | 0 | No Data
|  | fL | MCV |  |  |  | 0 | No Data
| High | x10⁹/L | WBC |  |  |  | 0 | No Data
|  | x10¹²/L | RBC |  |  |  | 0 | No Data
|  | g/L | HGB |  |  |  | 0 | No Data
|  | x10⁹/L | PLT |  |  |  | 0 | No Data
|  | % | HCT |  |  |  | 0 | No Data
|  | fL | MCV |  |  |  | 0 | No Data
Post Calibration
I. Coefficient
Lot Number: | [6 empty] | Expiration Date:
|  | Parameter |  | WBC | RBC | HGB | MCV | PLT
Coefficient (%)
|  | Coeff. Range (%) |  | 75 - 125 | 76 - 125 | 77 - 125 | 78 - 125 | 79 - 125
|  | Remarks |  | Fail | Fail | Fail | Fail | Fail
II. Control Data
| Level | Unit | Para. | Measured | Target | Range | Error | Remarks
| Low | x10⁹/L | WBC |  |  |  | 0 | No Data
|  | x10¹²/L | RBC |  |  |  | 0 | No Data
|  | g/L | HGB |  |  |  | 0 | No Data
|  | x10⁹/L | PLT |  |  |  | 0 | No Data
|  | % | HCT |  |  |  | 0 | No Data
|  | fL | MCV |  |  |  | 0 | No Data
| Normal | x10⁹/L | WBC |  |  |  | 0 | No Data
|  | x10¹²/L | RBC |  |  |  | 0 | No Data
|  | g/L | HGB |  |  |  | 0 | No Data
|  | x10⁹/L | PLT |  |  |  | 0 | No Data
|  | % | HCT |  |  |  | 0 | No Data
|  | fL | MCV |  |  |  | 0 | No Data
| High | x10⁹/L | WBC |  |  |  | 0 | No Data
|  | x10¹²/L | RBC |  |  |  | 0 | No Data
|  | g/L | HGB |  |  |  | 0 | No Data
|  | x10⁹/L | PLT |  |  |  | 0 | No Data
|  | % | HCT |  |  |  | 0 | No Data
|  | fL | MCV |  |  |  | 0 | No Data
Remarks
Performed by: | [7 empty] | Noted by:
Engr. Cris Ian H. Montero
Service Engineer | [7 empty] | Technical Manager
MECHANICAL DATA
Certificate Number:
Needle Position
Needle Position 1: |  |  |  | Steps |  | Needle Position 2: |  |  | Steps
Needle Position 3: |  |  |  | Steps |  | Needle Position 4: |  |  | Steps
Needle Position 5: |  |  |  | Steps
Pressure
Target Value (0 - 4000): | [5 empty] | Current Value:
Test Start Pressure: | [5 empty] | Test Max Pressure
Hemoglobin
HGB Blank AD: | [5 empty] | HGB Feedback:
Engineering Parameters
WBC Output: | [5 empty] | PLT Burst Lower Limit:
WBC Burst Lower Limit: | [5 empty] | PLT Burst Upper Limit:
RBC Output: | [5 empty] | PLT Output:
RBC Burst Lower Limit:
Pulse Width Parameter 1:
Pulse Width Parameter 2: | [5 empty] | Pulse Width Parameter 3:
Count Time (S): | [5 empty] | Pulse Width Parameter 4:
Warning Time of Clogging (S): | [5 empty] | Count Wait Time (S):
Warning Time of Air Bubble (S):
Others
Hardware alarm reagent amount of correction: | [5 empty] | ml
Remarks
Performed by: | [7 empty] | Noted by:
Engr. Cris Ian H. Montero
Service Engineer | [7 empty] | Technical Manager

## Sample Only - Do not Use
A L T A M E D I C A I N C O R P O R A T E D Medical Products You Can Trust!
CERTIFICATE OF CALIBRATION
Certificate Number: |  | 122235 | [6 empty] | 2020-03-02
Performed by:
Company Name: |  | Altamedica Incorporated
Address: |  | 2nd Floor Masonic Temple Bldg. J.M. Basa Street, Iloilo City
Contact Number: |  | Tel#: (033) 508 0208 / (033) 333 1769, Fax#: (033) 508 0208
Customer:
Company Name: |  | Medicus Philippines Incorporated - Molo
Address: |  | Locsin St., Barangay Tap-oc, Molo, Iloilo City, 5000 Iloilo City
Contact Number: |  | (033) 509 4590
Equipment:
Equipment: |  | 3-Part Hematology Analyzer |  |  |  | Manufacturer: |  | Linear Chemicals
Model: |  | STEL-3 |  |  |  | Manufacture Date: |  | 2018-04-30
Serial Number: |  | 802002007IEXN |  |  |  | Distributor: |  | Altamedica Incorporated
As received: |  | Out of Tolerance |  |  |  | As Released: |  | In Tolerance
Test Condition:
Room Temperature: |  | 25 |  |  |  | Relative Humidity: |  | 60%
Performed at: |  | Medicus Molo - Locsin St, Barangay Tap-oc, Molo, Iloilo City, 5000 Iloilo
Methods and Materials:
Method(s): |  | Tri-Level Control Calibration
Procedure(s): |  | Hematology Analyzer Work Instruction Manual
Material(s) used: |  | YX360TRF Sanwa Multimeter |  |  |  | Serial Number: |  | N/A
|  | Tri-Level Hematology Control |  |  |  | Brand: |  | Biorad Liquicheck Hematology-16
[6 empty] | Lot Number: |  | 77850
[6 empty] | Expiration Date: |  | 2020-04-12
Calibration Result:
Calibration Date: |  | 2020-02-05 |  |  |  | Calibration Due: |  | 2021-02-05
Background Test: |  | Pass |  | Repeatability: |  | Pass |  | Accuracy: |  | Pass
Conditions/Analysis: |  | Passed! Post calibration parameters are within acceptable limit.
This report may not be reproduced, except in full, without the written approval of the Technical Manager. The results reported herein apply only to the calibration of the item described above, and no limitations of use apply to the calibrated unit. Although the calibrated item meets the specifications and performance at the time of calibration, due to any number of factors, the recommended due date of the item calibrated does not imply continuing confidence to specifications during the recommended interval.
Performed by: | [7 empty] | Noted by:
Leonil C. Cagcon, ECE | [7 empty] | Engr. Cris Ian H. Montero
Service Engineer | [7 empty] | Technical Manager
ANALYTICAL DATA
Certificate Number: 122235
2020-03-02
PRE - CALIBRATION COEFFICIENT
|  | Parameter |  | WBC | RBC | HGB | MCV | PLT
|  | Coefficient (%) |  | 100 | 100 | 100 | 100 | 100
|  | Coeff. Range (%) |  | 75 - 125 | 76 - 125 | 77 - 125 | 78 - 125 | 79 - 125
|  | Remarks |  | Pass | Pass | Pass | Pass | Pass
PRE - CALIBRATION CONTROL DATA
| Level | Unit | Para. | Measured | Target | Range | Error | Remarks
| Low | x10⁹/L | WBC | 3.47 | 3.4 | 0.4 | 0.07 | Pass
|  | x10¹²/L | RBC | 2.29 | 2.32 | 0.15 | 0.03 | Pass
|  | g/L | HGB | 63 | 64 | 4 | 1 | Pass
|  | x10⁹/L | PLT | 58 | 49 | 18 | 9 | Pass
|  | % | HCT | 21 | 20 | 2.7 | 1 | Pass
|  | fL | MCV | 91.6 | 86.2 | 6 | 5.4 | Pass
| Normal | x10⁹/L | WBC | 8.95 | 8.7 | 1 | 0.25 | Pass
|  | x10¹²/L | RBC | 4.2 | 4.07 | 0.18 | 0.13 | Pass
|  | g/L | HGB | 123 | 124 | 5 | 1 | Pass
|  | x10⁹/L | PLT | 204 | 200 | 35 | 4 | Pass
|  | % | HCT | 41.3 | 38.9 | 3.3 | 2.4 | Pass
|  | fL | MCV | 98.4 | 95.6 | 4 | 2.8 | Pass
| High | x10⁹/L | WBC | 18.88 | 18.4 | 2 | 0.48 | Pass
|  | x10¹²/L | RBC | 5.06 | 4.91 | 0.2 | 0.15 | Pass
|  | g/L | HGB | 163 | 160 | 6 | 3 | Pass
|  | x10⁹/L | PLT | 431 | 406 | 60 | 25 | Pass
|  | % | HCT | 52.4 | 48.8 | 4 | 3.6 | Pass
|  | fL | MCV | 103.6 | 99.4 | 4 | 4.2 | Fail
BACKGROUND TEST
|  | Parameter |  | WBC | RBC | HGB | HCT | PLT
|  | Unit |  | x10⁹/L | x10¹²/L | g/L | % | x10⁹/L
|  | Acceptable Range |  | ≤0.2 | ≤0.02 | ≤1 | ≤0.5 | ≤10
|  | Background Test |  | 0.01 | 0 | 0 | 0 | 0
|  | Remarks |  | Pass | Pass | Pass | Pass | Pass
REPEATABILITY
|  | Parameter |  | WBC | RBC | HGB | MCV | PLT
|  | 1 |  | 8.98 | 4.15 | 123 | 98.36 | 202.05
|  | 2 |  | 8.86 | 4.16 | 123 | 98.69 | 207.97
|  | 3 |  | 9.05 | 4.14 | 124.04 | 98 | 206.08
|  | 4 |  | 8.79 | 4.11 | 122.66 | 98.43 | 204.86
|  | 5 |  | 8.96 | 4.12 | 122.68 | 98.43 | 193.79
|  | MEAN |  | 8.928 | 4.136 | 123.076 | 98.382 | 202.95
|  | SD |  | 0.10281 | 0.020736 | 0.56363 | 0.24793 | 5.5534
|  | CV (%) |  | 1.1516 | 0.50136 | 0.45795 | 0.25201 | 2.7363
|  | Allowed CV (%) |  | 2 | 1.5 | 1.5 | 0.4 | 4
|  | Remarks |  | Pass | Pass | Pass | Pass | Pass
POST - CALIBRATION COEFFICIENT
|  | Parameter |  | WBC | RBC | HGB | MCV | PLT
|  | Coefficient (%) |  | 97.45 | 98.4 | 100.62 | 97.25 | 97.58
|  | Remarks |  | Pass | Pass | Pass | Pass | Pass
POST - CALIBRATION QUALITY CONTROL DATA
| Level | Unit | Para. | Measured | Target | Range | Error | Remarks
| Low | x10⁹/L | WBC | 3.39 | 3.4 | 0.4 | 0.01 | Pass
|  | x10¹²/L | RBC | 2.26 | 2.32 | 0.15 | 0.06 | Pass
|  | g/L | HGB | 64 | 64 | 4 | 0 | Pass
|  | x10⁹/L | PLT | 54 | 49 | 18 | 5 | Pass
|  | % | HCT | 20.2 | 20 | 2.7 | 0.2 | Pass
|  | fL | MCV | 89 | 86.2 | 6 | 2.8 | Pass
| Normal | x10⁹/L | WBC | 8.73 | 8.7 | 1 | 0.03 | Pass
|  | x10¹²/L | RBC | 3.97 | 4.07 | 0.18 | 0.1 | Pass
|  | g/L | HGB | 124 | 124 | 5 | 0 | Pass
|  | x10⁹/L | PLT | 189 | 200 | 35 | 11 | Pass
|  | % | HCT | 38 | 38.9 | 3.3 | 0.9 | Pass
|  | fL | MCV | 95.6 | 95.6 | 4 | 0 | Pass
| High | x10⁹/L | WBC | 18.4 | 18.4 | 2 | 0 | Pass
|  | x10¹²/L | RBC | 4.98 | 4.91 | 0.2 | 0.07 | Pass
|  | g/L | HGB | 164 | 160 | 6 | 4 | Pass
|  | x10⁹/L | PLT | 400 | 406 | 60 | 6 | Pass
|  | % | HCT | 50.1 | 48.8 | 4 | 1.3 | Pass
|  | fL | MCV | 100.6 | 99.4 | 4 | 1.2 | Pass
Remarks
Passed! Post calibration parameters are within acceptable limit.
Performed by: | [7 empty] | Noted by:
Leonil C. Cagcon, ECE | [7 empty] | Engr. Cris Ian H. Montero
Service Engineer | [7 empty] | Technical Manager
MECHANICAL DATA
Certificate Number: 122235
2020-03-02
Needle Position
Needle Position 1: |  |  |  | Steps |  | Needle Position 2: |  |  |  | Steps
Needle Position 3: |  |  |  | Steps |  | Needle Position 4: |  |  |  | Steps
Needle Position 5: |  |  |  | Steps
Pressure
Target Value (0 - 4000): | [5 empty] | Current Value:
Test Start Pressure: |  |  | 2600 | 23 |  | Test Max Pressure |  |  | 2600 | 23
Hemoglobin
HGB Blank AD: | [5 empty] | HGB Feedback:
Engineering Parameters
WBC Output: | [5 empty] | PLT Burst Lower Limit:
WBC Burst Lower Limit: | [5 empty] | PLT Burst Upper Limit:
RBC Output: | [5 empty] | PLT Output:
RBC Burst Lower Limit:
Pulse Width Parameter 1:
Pulse Width Parameter 2: | [5 empty] | Pulse Width Parameter 3:
Count Time (S): | [5 empty] | Pulse Width Parameter 4:
Warning Time of Clogging (S): | [5 empty] | Count Wait Time (S):
Warning Time of Air Bubble (S):
Others
Hardware alarm reagent amount of correction: | [5 empty] | ml
Remarks
.
Performed by: | [7 empty] | Noted by:
Leonil C. Cagcon, ECE | [7 empty] | Engr. Cris Ian H. Montero
Service Engineer | [7 empty] | Technical Manager