## Blank Template
A L T A M E D I C A I N C O R P O R A T E D Medical Products You Can Trust!
CERTIFICATE OF CALIBRATION
Certificate Number: | [7 empty] | Date:
Performed by:
Company Name: |  | Altamedica Incorporated
Address: |  | 2nd Floor Masonic Temple Bldg. J.M. Basa Street, Iloilo City
Contact Number: |  | Tel#: (033) 508 0208 / (033) 333 1769, Fax#: (033) 508 0208
Customer:
Company Name:
Address:
Contact Number:
Equipment:
Equipment: | [5 empty] | Manufacturer:
Model: | [5 empty] | Manufacture Date:
Serial Number: | [5 empty] | Distributor:
As received: | [5 empty] | As Released:
Test Condition:
Room Temperature: | [5 empty] | Relative Humidity:
Performed at:
Methods and Materials:
Method(s): |  | Tri-Level Control Calibration
Procedure(s): |  | Hematology Analyzer Work Instruction Manual
Material(s) used: |  | Description |  |  |  | Serial Number / Lot Number |  |  | Expiry
|  | R&D Hematology Control - Low |  |  |  | 112345
|  | R&D Hematology Control - Normal |  |  |  | 112346
|  | R&D Hematology Control - High |  |  |  | 112347
Calibration Result:
Calibration Date: | [5 empty] | Calibration Due:
Background Test: |  |  |  | Repeatability: |  |  |  | Accuracy:
Conditions/Analysis:
This report may not be reproduced, except in full, without the written approval of the Technical Manager. The results reported herein apply only to the calibration of the item described above, and no limitations of use apply to the calibrated unit. Although the calibrated item meets the specifications and performance at the time of calibration, due to any number of factors, the recommended due date of the item calibrated does not imply continuing confidence to specifications during the recommended interval.
Performed by: | [7 empty] | Noted by:
Engr. Cris Ian H. Montero
Service Engineer | [7 empty] | Technical Manager
ANALYTICAL DATA
Certificate Number:
Background Test
|  | Parameter |  | WBC | RBC | HGB | HCT | PLT
|  | Unit |  | x10⁹/L | x10¹²/L | g/L | % | x10⁹/L
|  | Acceptable Range |  | 0.2 | 0.02 | 1 | 0.5 | 5
Background Test
|  | Remarks |  | Pass | Pass | Pass | Pass | Pass
Repeatability
|  | Parameter |  | WBC | RBC | HGB | MCV | PLT
|  | 1 |  | 5.71 | 4.43 | 131 | 95.5 | 318
|  | 2 |  | 5.77 | 4.4 | 130 | 95.4 | 318
|  | 3 |  | 5.52 | 4.4 | 130 | 95.9 | 334
|  | 4 |  | 5.66 | 4.36 | 128 | 95.8 | 330
|  | 5 |  | 5.65 | 4.39 | 130 | 96.1 | 320
|  | MEAN |  | 5.662 | 4.396 | 129.8 | 95.74 | 324
|  | SD |  | 0.092574 | 0.0251 | 1.0954 | 0.2881 | 7.4833
|  | CV (%) |  | 1.635 | 0.57097 | 0.84395 | 0.30092 | 2.3097
|  | Allowed CV (%) |  | 2 | 1.5 | 1.5 | 1 | 4
|  | Remarks |  | Pass | Pass | Pass | Pass | Pass
PRE - CALIBRATION
I. Coefficient
|  | Parameter |  | WBC | RBC | HGB | MCV | PLT
Coefficient (%)
|  | Coeff. Range (%) |  | 75 - 125 | 76 - 125 | 77 - 125 | 78 - 125 | 79 - 125
|  | Remarks |  | Fail | Fail | Fail | Fail | Fail
II. Control Data
Lot Number: | [6 empty] | Expiration Date:
| Level | Unit | Para. | Measured | Target | Range | Error | Remarks
| Low | x10⁹/L | WBC |  |  |  | 0 | No Data
|  | x10¹²/L | RBC |  |  |  | 0 | No Data
|  | g/L | HGB |  |  |  | 0 | No Data
|  | x10⁹/L | PLT |  |  |  | 0 | No Data
|  | % | HCT |  |  |  | 0 | No Data
|  | fL | MCV |  |  |  | 0 | No Data
| Normal | x10⁹/L | WBC |  |  |  | 0 | No Data
|  | x10¹²/L | RBC |  |  |  | 0 | No Data
|  | g/L | HGB |  |  |  | 0 | No Data
|  | x10⁹/L | PLT |  |  |  | 0 | No Data
|  | % | HCT |  |  |  | 0 | No Data
|  | fL | MCV |  |  |  | 0 | No Data
| High | x10⁹/L | WBC |  |  |  | 0 | No Data
|  | x10¹²/L | RBC |  |  |  | 0 | No Data
|  | g/L | HGB |  |  |  | 0 | No Data
|  | x10⁹/L | PLT |  |  |  | 0 | No Data
|  | % | HCT |  |  |  | 0 | No Data
|  | fL | MCV |  |  |  | 0 | No Data
POST - CALIBRATION
I. Coefficient
|  | Parameter |  | WBC | RBC | HGB | MCV | PLT
Coefficient (%)
|  | Coeff. Range (%) |  | 75 - 125 | 76 - 125 | 77 - 125 | 78 - 125 | 79 - 125
|  | Remarks |  | Fail | Fail | Fail | Fail | Fail
II. Control Data
Lot Number: | [6 empty] | Expiration Date:
| Level | Unit | Para. | Measured | Target | Range | Error | Remarks
| Low | x10⁹/L | WBC |  |  |  | 0 | No Data
|  | x10¹²/L | RBC |  |  |  | 0 | No Data
|  | g/L | HGB |  |  |  | 0 | No Data
|  | x10⁹/L | PLT |  |  |  | 0 | No Data
|  | % | HCT |  |  |  | 0 | No Data
|  | fL | MCV |  |  |  | 0 | No Data
| Normal | x10⁹/L | WBC |  |  |  | 0 | No Data
|  | x10¹²/L | RBC |  |  |  | 0 | No Data
|  | g/L | HGB |  |  |  | 0 | No Data
|  | x10⁹/L | PLT |  |  |  | 0 | No Data
|  | % | HCT |  |  |  | 0 | No Data
|  | fL | MCV |  |  |  | 0 | No Data
| High | x10⁹/L | WBC |  |  |  | 0 | No Data
|  | x10¹²/L | RBC |  |  |  | 0 | No Data
|  | g/L | HGB |  |  |  | 0 | No Data
|  | x10⁹/L | PLT |  |  |  | 0 | No Data
|  | % | HCT |  |  |  | 0 | No Data
|  | fL | MCV |  |  |  | 0 | No Data
Remarks
Performed by: | [7 empty] | Noted by:
Engr. Cris Ian H. Montero
Service Engineer | [7 empty] | Technical Manager
MECHANICAL DATA
Certificate Number:
Touch Screen
Perform Touch Screen Calibration?
Remarks:
Vaccuum Tank
Vaccuum Pressure (-33 ~ -25): |  |  |  |  | Fail
Reading Chamber
Vaccuum Pressure (-33 ~ -25): |  |  |  |  | Fail |  | Diluent Voltage (3 ~ 5): |  |  | Fail
WBC Aperture Voltage (1.2 ~ 2.-5): |  |  |  |  | Fail |  | Lyse Voltage (3 ~ 5) |  |  | Fail
RBC Aperture Voltage(1.2 ~ 2.-5): |  |  |  |  | Fail
Hemoglobin
HGB Gain (4.2 ~ 4.8): |  |  |  |  | Fail
Power Supply
A+ 12V Power Supply (10.8 ~ 11.2) |  |  |  |  | Fail
A- 12V Power Supply (-13.2 ~ -10.8) |  |  |  |  | Fail
Remarks
Performed by: | [7 empty] | Noted by:
Engr. Cris Ian H. Montero
Service Engineer | [7 empty] | Technical Manager