INTRODUCTION

Thank you for your purchase of Z3/Z3 CRP  Series Hematology Analyzer

manufactured by Zybio Inc.

(Hereinafter referred to as "Zybio")

Before using the product,

please carefully read the Operation Manual  so as to use the product correctly.

Please retain this operation manual after reading,  so that you can consult it as needed.

Product name Hematology AnalyzerProduct models Z3 CRP, Z3, Z31

Product features and  composition

Intended use

It consists of a main unit and accessories, including an automatic sampling  system and a data calculation and processing system.

The product works in combination with the matched reagents produced by  Zybio for quantitative detection of the analytes in human blood samples in  clinical settings through the electric impedance, colorimetric and scattering  turbidimetry methods.

The 23 test items include white blood cell count (WBC #), lymphocyte count  (Lym #), intermediate cell count (Mid #), neutrophil count (Gran #), lymphocyte  percentage (Lym %), intermediate cell percentage (Mid %), neutrophil percentage  (Gran %), red blood cell count (RBC #), hemoglobin (HGB), mean corpuscular  volume (MCV), mean corpuscular hemoglobin (MCH), mean corpuscular  hemoglobin concentration (MCHC), red blood cell distribution width-coefficient  of variation (RDW-CV), red blood cell distribution width-standard deviation  (RDW-SD), hematocrit (HCT), platelet count (PLT #), mean platelet volume  (MPV), platelet distribution width (PDW), plateletcrit (PCT), platelet-large cell  ratio (P-LCR), platelet-large cell count (P-LCC), C-reactive protein (CRP), and  hypersensitive C-reactive protein (Hs-CRP). Models z3 and z31 do not test  C-reactive protein (CRP) or hypersensitive C-reactive protein (hs-CRP). Number of product technical requirements / Registration certificate number:  YXZZ 20182400068

Production license number: YSYJXSCX 20150016

Manufacturer Zybio Inc.

Domicile of registrant Floor 1 to Floor 4, Building 30, No.6 of Taikang Road, Block C of Jianqiao  Industrial Park, Dadukou District, Chongqing, China 400082

Production address Floor 1 to Floor 4, Building 30, No.6 of Taikang Road, Block C of Jianqiao  Industrial Park, Dadukou District, Chongqing, China 400082

Date of production See the nameplate of the main unit

L

A

U

N

A

Service life 7 years. This service life is determined according to the lifespan test performed

M

1

S

I

on the instrument. In the course of use, the user shall carry out maintenance

H

R

T

E

T

and repair of the product according to the Operation Manual. After maintenance

G

P

N

A

I

and repair, a product that has been confirmed to maintain its basic safety and

H

S

C

U

effectiveness can be used normally.

1

0

0

Release date 2017-10-12

Intellectual

Property Statement Statement

Zybio Inc. (hereinafter referred to as “Zybio”) owns the intellectual property rights  to this operation manual. This manual may refer to information protected by  copyright or patents and does not convey any license under the patent rights or  copyright of Zybio, or of others. The intellectual property rights of the MANUAL  and the corresponding product belong to Zybio Inc.

Zybio intends to maintain the contents of this manual as confidential information.  Disclosure of the information in this manual in any manner whatsoever without the  written permission of Zybio is strictly forbidden.

Release, amendment, reproduction, distribution, rental, adaptation, translation or  any other derivative work of this manual in any manner whatsoever without the  written permission of Zybio is strictly forbidden.

Zybio is a registered trademark or trademark of Zybio in China and other  countries . All other trademarks that appear in this manual are used only for  informational or editorial purposes. They are the property of their respective  owners.

Zybio reserves the final right to interpret this manual.

In the event that all of the following requirements are met, Zybio is  responsible for the safety, reliability and performance of the product: • The assembly, extensions, readjustments, modifications, and repairs are  performed by professionals authorized by Zybio.

• All replacement parts used in the repairs and all accessories and consumables  used are products of or approved by Zybio.

• Relevant electrical equipment conforms to national standards and the  requirements of this manual.

• The operation of the product shall be carried out in accordance with this manual.

Exemptions

After-sales service Warning

Zybio’s obligation or liability under this warranty does not include any  transportation or other charges or liability for direct, indirect or consequential  damages or delay resulting from the improper use or application of the product or  the use of parts or accessories not approved by Zybio or repairs by people other  than Zybio authorized personnel.

This warranty shall not extend to:

• Malfunction or damage caused by improper use or man-made failure. • Malfunction or damage caused by unstable or out-of-range power input. • Malfunction or damage caused by force majeure such as fire and earthquake. • Malfunction or damage caused by improper operation or repair by unqualified or  unauthorized service people.

• Malfunction of the instrument or part whose serial number is not legible enough. • Others not caused by instrument or part itself.

Contact: Zybio Inc.

Address: Floor 1 to Floor 4, Building 30, No.6 of Taikang Road, Block C of  Jianqiao Industrial Park, Dadukou District, Chongqing, China, 400082 Te l: +86 (0)23-6865 5509 Website: www.zybio.com  Fax: +86 (0)23-6869 9779 24-hour service hotline: 400-056-7879

EC- Representative: Shanghai International Holding Corp. GmbH(Europe) Address: Eiffestraβe 80, 20537 Hamburg, Germany

Te l: 0049-40-2513175 Fax: 0049-40-255726

• This system can only be operated by laboratory professionals, doctors, or  laboratory technicians who have been trained by Zybio or by the agents of Zybio. • If the hospitals or institutions responsible for using this instrument fail to  implement a satisfactory maintenance/repair plan, it may cause abnormal

1

R

ET

PA

HC

Warranty THIS WARRANTY IS EXCLUSIVE AND IS IN LIEU OF ALL OTHER

WARRANTIES, EXPRESSED OR IMPLIED, INCLUDING WARRANTIES OF

MERCHANTABILITY OF FITNESS FOR ANY PARTICULAR PURPOSE.

L

A

U

N

A

M

S

I

H

T

G

N

I

S

U

Precautions

instrument failure and may endanger personal health.

• Ensure that the analyzer is used under the conditions specified in the MANUAL.  If the usage conditions are not met, the analyzer may not operate normally, its

L

measurement results will not be reliable, or its components may be damaged and

A

U

N

personal safety may be endangered.

A

M

1

S

I

H

R

T

E

T

G

P

N

A

I

• This equipment must be used by skilled/trained clinical professionals.

H

S

C

U

2 00

3 00

CONTENTS

AAAAAAAA CHAPTER 4 CHAPTER 8 APPENDIX A

001 002

Introduction

Installing Your Analyzer Calibrating Your Analyzer Specifications 032 072 102

002 002

Intellectual Property StatementI Statement

Warranty and Exemptions

Introduction

032 072

Installation Requirements

035 073 Connecting the Analyzer System

Introduction

When to Calibrate How to Calibrate

102 102

Classification Reagents

Applicable Tubes

003 003 003

After-sales service Waring

Precautions

038 039

Installing the thermal paper Notes

CHAPTER 9

Customizing the Analyzer Software

103 103 105 106

Parameters

Sampling Features

Input and Output Devices Main Unit Interfaces

CHAPTER 1

CHAPTER 5

Operating Your Analyzer

082 083

Introduction

General Settings

106 106 106

Power Supply

Fuse

Electromagnetic Compatibility

Using This Manual

040

Introduction

086

System Settings6

107

Sound Pressure9

006 006 006 007

Introduction

Who Should Read This Manual How to Find Information

Conventions used in This Manual

041 042 043 043

Initial Checks

Startup and Login Daily Quality Control Sample Preparation

088 089

Parameter Settings User Settings

107 107 107 107

Operating Environment Dimensions and Weights Contraindications

Safety Classification

008 008

Common Operations Safety Information

045 052 052

Sample Analysis Sleep

Shutdown

CHAPTER 10

Servicing Your Analyzer

APPENDIX B

CHAPTER 2

Understanding Your Analyzer

CHAPTER 6

092 093 094 095

Introduction

Maintenance

Self-check

System Calibration

108

Key Parts

010 010

Introduction

Product Models

Reviewing Sample Results

095 096

Log

APPENDIX C

010 014 015

Product Description

Operating Interface

Reagents, Controls and Calibrators

Introduction Status

054

109

054

Sample review

CHAPTER 11

List of Spare Parts

CHAPTER 7 APPENDIX D

Troubleshooting Your Analyzer

110

CHAPTER 3

Using the QC ProgramsToxic and Hazardous Substances

1

L

A

U

N

Understanding the System Principles

A

M

S

I

060 061

Introduction L-J QC

098 098 101

Introduction

Error Information and Handling Repair

or Elements

L

A

U

N

A

M

1

S

I

R

ET

PA

HC

H

T

G

NI

SU

018 018 018

Introduction Aspiration Dilution

069

X-B QC

R

ET

PA

HC

H

T

G

NI

SU

4 00

025 028 031 031

WBC/HGB Measurement RBC/PLT Measurement CRP Measurement

Rinse

5 00

CHAPTER 1

USING THIS MANUAL

CHAPTER INTRODUCTIONW

Using the QC Programs Introduces the basic requirements for QC and the QC methods  provided by Z3/Z3 CRP Series Hematology Analyzer.

Introduction

This chapter explains how to use your Z3/Z3 CRP Series’ operation manual,  which is shipped with your Z3/Z3 CRP Series HEMATOLOGY ANALYZER and

Using the Calibration  Programs

Customizing the Analyzer  Software

Introduces the basic requirements for calibration and the calibration  methods provided by Z3/Z3 CRP Series Hematology Analyzer.

Introduces the configuration of system parameters, such as  software date format and parameter units.

Who Should Read  This Manual

contains reference information about the analyzer and procedures for operating,  troubleshooting and maintaining the analyzer. Read this manual carefully before  operating your Z3/Z3 CRP Series analyzer and operate your Z3/Z3 CRP Series  analyzer strictly as instructed in this manual.

This manual is intended to be read by clinical laboratory professionals. This  equipment must only be operated by skilled/trained clinical professionals. This  information contains information for clinical laboratory professionals to: • Learn about the hardware and software of Z3/Z3 CRP Series Hematology Analyzer;  • Customize system parameters;

Maintaining Your Analyzer Introduces the maintenance and testing processes of Z3/Z3 CRP  Series Hematology Analyzer.

Troubleshooting Your Analyzer Introduces the troubleshooting processes of Z3/Z3 CRP Series  Hematology Analyzer.

Appendix A Specifications Introduces the specifications of Z3/Z3 CRP Series Hematology  Analyzer.

Appendix B Key Parts Introduces the key parts of Z3/Z3 CRP Series Hematology  Analyzer.

Appendix C List of Spare

Parts Lists the spare parts for Z3/Z3 CRP Series Hematology Analyzer. Appendix D Names and

• Perform daily operations;

• Perform system maintenance and troubleshooting.

Contents of Toxic and  Hazardous Substances or  Elements

Introduces the toxic and harmful substances or elements in the  Z3/Z3 CRP Series Hematology Analyzer and their contents.

1

R

ET

How to Find

Information

L

A

U

N

A

M

S

I

H

T

G

This operation manual comprises 11 chapters and 4 appendices. Refer to the  table below to find the information you need.

CHAPTER INTRO DUCTION

Using This Manual Introduces how to use this manual of Z3/Z3 CRP Series  Hematology Analyzer.

Understanding Your Analyzer Introduces the composition, software interface, and software  operation of Z3/Z3 CRP Series Hematology Analyzer.

Working Principles Introduces the measuring principles and workflow of the  Z3/Z3 CRP Series Hematology Analyzer.

Installing Your Analyzer Introduces the installation requirements and installation methods  for Z3/Z3 CRP Series Hematology Analyzer.

Introduces daily operations such as the methods of sample

Conventions used  in This Manual

This manual uses different fonts and formats to distinguish content with special  meanings in the text.

FORMAT DESCRIPTION

[××] ×× stands for a button in an external keyboard or panel

“××” ×× stands for the information displayed in the interface

×× ×× stands for the quoted chapter

L

A

All illustrations provided in this manual should be used only for reference. The

U

N

A

graphs, settings, or data in the illustrations may not exactly match the actual

M

1

S

I

display you see on Z3/Z3 CRP Series Hematology Analyzer.

H

R

T

E

T

G

PA

HC

NI

SU

Operating Your Analyzer

collection and preparation, the process of sample analysis,  and turning the system on/off.

PA

HC

NI

SU

6

7

Reviewing Sample Results Introduces the process for reviewing of the results of the sample

0

0

analysis.

0

0

Common Operations ACTION OPERATION PERFORMED

Click Tap the ×× key or button with your finger or click on ×× using the

left mouse button.

Click the “××” edit box to move the cursor to the appropriate field

The analyzer system may contain the following symbols: WHEN YOU SEE… MEANING

Biological risk

Enter

Delete

Select from the drop-down list  ×× (only for a drop-down list)

and use the keyboard or on-screen keyboard to complete the  data entry

Click the left mouse button, or tap directly on the touch screen,  or use the [←][→][Home][End] keys on the external keyboard to  move the cursor to the point where you want to delete, and then  use the [Del] key to delete the character following the cursor or  use the [BackSpace] key (the [←] key in the upper right corner of  the on-screen keyboard) to delete the character before the cursor.

Click on the down arrow button in the “××” box to bring up the  drop-down list, (drag the scroll bar to) browse through the current  list, and then click on the field in the current list to select; or use  the [↑][↓] [PageUp][PageDown] keys to browse the current list and  press [Enter] to select the field where the arrow is located.

IVD LOT

Caution, Consult accompanying Documents Exercise caution when working Around to aviod pricking Protective earth (ground)

Alternating current

For in vitro diagnostic use

Batch code

1

R

ET

PA

HC

Safety Information Symbols used in the MANUAL:

WHEN YOU SEE… MEANING

read the statement below the symbol. The statement is alerting

you

to a potentially biohazardous condition.

read the statement below the symbol. The statement is alerting

you

to an operating hazard that can cause personnel injury.

read the statement below the symbol. The statement is alerting

you

to a possibility of analyzer damage or unreliable analysis results.

read the statement below the symbol. The statement is alerting

you

to information that requires your attention.

L

A

U

N

A

M

S

I

H

T

G

N

I

S

U

Use by

Serial number

Date of manufacture

Manufacturer

Consult the operation manual

Temperature limitation

The device is fully conformance

With the council directive

L

Concerning in vitro diagnostic

A

Medical devices 98/79/ec.

U

N

A

M

1

S

I

H

R

T

E

T

G

P

N

A

I

H

S

C

U

8 00

9 00

Introduction Product Models

CHAPTER 2

UNDERSTANDING YOUR  ANALYZER

Z3/Z3 CRP SERIES HEMATOLOGY ANALYZER works in combination with the  diagnostic reagents produced by Zybio for quantitative detection of the analytes  in human blood samples in clinical settings through electric the resistance  method, colorimetric method and scattering turbidimetry method.

This product has three models: Z3 CRP, Z3, and Z31. The working principles,  main functions, electrical structure and key components of the various models

1

2

3

4

Figure 2-1 Front of the Main Unit

are basically the same. The only difference is their functional configuration (see  Table 1 below for details). Therefore, the full-featured Z3 CRP model can cover  the Z3 and Z31 models during the registration check.

1-Touch screen 2-Sample probe

3-Aspirate key 4-Indicator light

MODEL FUNCTIONAL CONFIGURATION

Z3 CRP Contains the full functionality of all models, with a testing throughput of 70  samples per hour

Compared with the Z3 CRP model, this model does not include the function of

This manual uses different fonts and formats to distinguish content with special  meanings in the text.

Z3

Z31

R

E

Z

Y

L

A

N

A

measuring the results of CRP-related parameters, and the testing throughput

is 70 samples per hour.

Compared with the Z3 CRP model, this model does not include the function

of measuring the results of CRP-related parameters or the function of providing

the results of PLCC/PLCR parameters, and the testing throughput is 60

samples per hour.

Table 2-1 Model Differences

1

R

E

4

Z

Y

L

A

2

N

A

5

2

R

ET

PA

HC

RU

O

Y

G

NI

D

N

AT

S

RE

D

NU

Product Description

Z3/Z3 CRP Series Hematology Analyzer comprises of a main unit and accessories,

R

U

O

6

including an automatic sampling system and a data calculation and processing

Y

G

3

system.

N

I

D

2

N

A

T

R

S

E

R

T

E

P

Figure 2-2 Backof the Main Unit

A

D

H

N

C

U

0 10

1-Waste sensor

2-Waste tube connector 3-Diluent tube connector

4-USB port

5-Network port

6-Power component

1 10

2

1

3

4

Figure 2-3 View with the Front Panels Removed

1

2

3

Figure 2-5 Left of the main unit

1-Syringe component 2-Sampling component

1

3-Aspirate key 4-Aspirating switch

1-Recorder

2-Side door lock

Touch Screen

3-Side door

2

RE

R EZ

YL

A

N

A

RU

O

Y

G

NI

D

N

AT

S

2

6

3

7

4

8

5

Figure 2-4 Right of the main unit (the right door opened)

The touch screen is located on the front of the analyzer and is used to perform  interface operations.

Aspirate Key

The aspirating key is located behind the sample probe and is used to start the  counting operation, add the diluent or draw the maintenance reagents.

Indicator Light

R

E

The indicator light is located on the front side and is used to indicate the current

Z

Y

L

A

status of the system, providing red, yellow and blue status indicators.

N

A

R

U

O

USB port

Y

G

The analyzer has 4 USB ports for connecting the external mouse, keyboard and

N

I

D

USB flash disk during debugging, maintenance and upgrading.

2

N

A

T

R

S

E

T

PA

HC

RE

D

NU

1-Air pump

2-Sampling component 3-Valve

5-R2 kit

6-Vacuum chamber 7-Liquid pump

R

T

E

P

Network Port

A

D

H

N

C

U

The analyzer has a network port on the back, which is used for connecting to an

2 10

4-CRP measuringcomponent

8-Counting chamber component

external computer to transmit data.

3 10

2

R

ET

PA

HC

Operating Interface Screen Display

After the startup completed, the "Sample Analysis" interface is displayed, as

shown in the following figure.

Figure 2-6 "Analysis" Interface

Menu Functions

Press the “Management" button to open the system menu as shown in the

following figure.

R

E

Z

Y

L

A

N

A

R

U

O

Y

G

N

I

D

N

A

T

S

R

E

D

N

U

Figure 2-7 “Management" Interface

Reagents

Controls and Calibrators

The analyzer, reagents, controls and calibrators together constitute a system,  which must be used as a whole to ensure proper performance. The operator  must use the reagents produced by Zybio. Otherwise, the analyzer may be  damaged and cannot meet the performance specifications described in the  MANUAL. Unspecified reagents cannot guarantee reliable analysis results.  “Reagents" in the MANUAL refer to the reagents used in combination with this  analyzer.

Before each reagent is used, the package must be checked. Damage to the  package may affect the quality of the reagent. Check the package for signs of  moisture or leakage. If such signs are noted, do not use the reagent.

• Please refer to the MANUAL of each reagent to use or store it.

• The operator should carry out background tests after replacing the diluent,  hemolytic agent (hereinafter referred to as “lyse”) or cleanser to ensure that the  background value is within the normal range, so as to prepare for sample analysis. • Ensure that the reagent is used before the expired date indicated on the  reagent’s labels.

• The reagents should stand motionless for a period of time until they become stable.

Reagents

• Diluent

This is an isotonic liquid with a specific electric conductivity. It is used to dilute   blood samples and provides a stable environment for blood cell counting. • Lyse

This is used for dissolving RBCs and generating hemoglobin complex. It is used   for human WBC counting, WBC grouping and hemoglobin determination. • R1 Reagent

This is a CRP reaction buffer used with R2 reagent to determine the content of   CRP in human serum.

R

• R2 Reagent

E

Z

Y

This is a CRP antibody used with R1 reagent to determine the content of CRP

L

A

N

in human serum.

A

R

• Probe Cleanser

U

O

Y

This is used for regular maintenance and cleaning of the analyzer.

G

N

I

D

2

N

A

T

R

S

E

R

T

E

P

A

D

H

N

C

U

4

5

1

1

The management menu includes 4 options, and the operator accesses the

0

0

corresponding interface through the system menu options to perform various

functions of the analyzer.

2

R

ET

PA

HC

Controls and Calibrators

The controls and calibrators are used for QC and calibration of the analyzer.

The controls are mainly composed of leucocyte-like cells, human erythrocytes,

platelet-like cells, preservatives and antiseptics. They are used for daily testing

WBC, RBC, HGB, MCV/HCT, PLT, and other parameters of Zybio’s analyzer,

so as to monitor or evaluate the precision of the results of the analyzer. There

are three levels of controls: low, normal and high. Daily QC runs can monitor the

operation of the analyzer to ensure the reliability of the results.

The calibrators are mainly composed of leucocyte-like cells, human erythrocytes,

platelet-like cells, preservatives and antiseptics, and they are used for calibrating

WBC, RBC, HGB, MCV / HCT, PLT and other parameters of Zybio’s Automated

Hematology Analyzer, thus establishing the metrological traceability of the results

of the analyzer.

Refer to the MANUALs of the controls and calibrators for their use and storage.

CRP Calibrators and CRP QC

CRP calibrators and CRP controls are used for the calibration and QC of the

analyzer.

CRP calibrators are an industrially-produced serum-like product used for the

calibration of this analyzer, thus establishing the metrological traceability of

CRP results of the analyzer. There are six calibrator levels: a, b, c, d, e, and f.

CRP controls are an industrially-produced serum-like product, which are used

to monitor and evaluate the precision of CRP results of the analyzer. There are

two levels of controls: I and II. Running the two levels of QC daily can monitor

the operation of the analyzer and ensure the reliability of the results. Refer to the

MANUAL of CRP controls and calibrators for their use and storage.

The "C-reactive protein (CRP) controls" and "C-reactive protein (CRP) calibrators"

mentioned in the MANUAL refer to the special controls and calibrators designated

R

R

E

E

Z

Z

by Zybio, and users must purchase them from Zybio or its agents.

Y

Y

L

L

A

A

N

N

A

A

R

R

U

U

O

O

Y

Y

G

G

N

N

I

I

D

D

2

N

N

A

A

T

T

R

S

S

E

R

R

T

E

E

P

D

A

D

N

H

N

U

C

U

6 10

7 10

3

R

ET

PA

HC

Introduction

Aspiration

Dilution

S

E

L

P

I

C

N

I

R

P

M

E

T

S

Y

S

E

H

T

G

N

I

D

N

A

T

S

R

E

D

N

U

CHAPTER 3

UNDERSTANDING THE  SYSTEM PRINCIPLES

The analyzer uses electrical impedance method to detect the number and volume  distribution of WBCs, RBCs and platelets; it uses colorimetry to measure the  hemoglobin concentration, and uses immunoturbidimetry to measure the CRP  content. Then, the analyzer calculates the results of other parameters.

In the whole blood (hereinafter referred to as WB) mode, the operator can send  the WB sample directly to the analyzer for sampling. In this case, the analyzer will  aspirate a quantified volume of WB sample.

In the pre-dilution (hereinafter referred to as PD) mode, the operator should first  mix 20 μL of capillary blood sample and 0.58 mL of diluent outside the machine  to form a diluted sample with a dilution ratio of 1: 30, and then send the diluted

sample to the analyzer for sampling. In this case, the analyzer will aspirate a  quantified volume of the diluted sample.

Various cells usually overlap each other in the samples submitted for testing.  In this case, the analyzer cannot accurately count blood cells or determine the  volume distribution of blood cells.

Therefore, the samples need to be diluted before the analyzer counts blood cells  or determines their volume distribution. The diluted sample allows single blood  cells to pass through the detecting aperture individually, and at the same time, it  provides a conductive environment for counting so as to facilitate cell counting  and determination of cell volume. Usually, RBCs are 1,000 times more than  WBCs.

Therefore, the interference of RBCs must first be eliminated before measuring  WBCs. RBCs usually have no nucleus, and the lyse can dissolve the erythrocyte  membrane so as to eliminate the RBCs in the WBC dilution.

Therefore, the lyse should be added to the diluted WBC samples to eliminate  RBCs before counting. The analyzer provides two different working modes - the

CBC WB mode

Draw sample 10μl

Add Diluent

Draw the second sample

Diluention sample

Add lyse Add Diluent

WBC sample RBC sample

Figure 3-1 The dilution process for WB mode

As shown in Figure 3-1, when the aspirated sample volume is 10 μL in the WB  mode, the analyzer aspirates a 10 μL WB sample which is then mixed with  the diluent to form a diluted sample. This diluted sample will be used in two  parts: One part is mixed with the diluent to form a secondary diluted sample,  which is used for counting RBCs and platelets and generating a corresponding

distribution histogram. The other part is mixed with the lyse to form another  S

E

sample, which is used to measure the hemoglobin concentration and count

L

P

I

C

WBCs, and generate a histogram of WBC distribution.

N

I

R

P

M

E

T

S

Y

S

E

H

T

G

N

I

D

3

N

A

T

R

S

E

R

T

E

P

A

D

H

N

C

U

WB mode and the PD mode for different types of samples.

8

9

1

1

0

0

The analyzer also provides three measurement modes - CBC, CBC + CRP,

and CRP.

CBC + CRP WB Mode CRP WB Mode

3

R

ET

PA

HC

Draw sample 15μl

Add R1 Add R2

Diluention sample CRP Sample

Add Diluent

Draw the second sample

Diluention sample

Add lyse Add diluent

WBC sample RBC sample

Figure 3-2 The dilution process for CBC + CRP WB mode

As shown in Figure 3-2, when the aspirated sample volume is 15 μL in the

CBC+CRP WB mode, the analyzer aspirates a 15 μL WB sample and 4 μL of the

S

E

sample is mixed with R1 to form a CRP diluted sample, and then R2 is added.

L

P

I

C

The resulting sample is used for CRP measurement.

N

I

R

P

The remaining sample is mixed with the diluent to form a diluted CBC sample.

M

E

T

This diluted sample will be used in two parts: A part is mixed with the diluent to

S

Y

S

form a secondary diluted sample, which is used for counting RBCs and platelets

E

H

T

and generating a corresponding distribution histogram. The remaining sample

G

N

I

is mixed with the lyse to form another sample, which is used to measure the

D

N

A

T

hemoglobin concentration and count WBCs, and generate a histogram of WBC

S

R

E

distribution.

D

N

U

Draw sample 15μl

Add R1

Diluention sample

Add R2

CRP sample

Figure 3-3 The dilution process for CRP WB mode

As shown in the figure above, when the aspirated sample volume is 15 μL in  the CRP WB mode, the analyzer aspirates a 15 μL WB sample and 4 μL of the  sample is mixed with R1 to form a CRP diluted sample, and then R2 is added.  One part is used for CRP measurement. The other part is tested for the HCT  value, which is used to correct the CRP results.

S

E

L

P

I

C

N

I

R

P

M

E

T

S

Y

S

E

H

T

G

N

I

D

3

N

A

T

R

S

E

R

T

E

P

A

D

H

N

C

U

0 20

1 20

CBC PD Mode

Peripheral blood

sample 20μl

A sample with a

diluention ratio of 1:30

Add Diluent 580μl

CBC + CRP PD Mode

Peripheral blood

sample 20μl

A sample with a

diluention ratio of 1:30

Add Diluent 580μl

The system sucks

183μl sample

Add Diluent

The system sucks

243μl sample

Diluention sample

Add Diluent

Add R2

CRP sample

Diluention sample

Draw the second sample

Add lyse Add Diluent

Diluention sample

Draw the second sample

Add lyse Add Diluent

3

R

ET

PA

HC

WBC sample RBC sample

Figure 3-4 The dilution process for PD mode

S

E

L

As shown in the figure above, in the PD mode, 20 μL of capillary blood is mixed

P

I

C

N

with the diluent (Out of Machine dilution) to form a diluted sample. The analyzer

I

R

P

aspirates 183 μL of the diluted sample and then it is added with the diluent to

M

E

T

S

form a secondary diluted sample. This diluted sample will be used in two parts:

Y

S

E

One part is mixed with the diluent to form a secondary diluted sample, which is

H

T

used for counting RBCs and platelets and generating a corresponding distribution

G

N

I

D

histogram. The other part is mixed with the lyse to form another sample, which is

N

A

T

S

used to measure the hemoglobin concentration and count WBCs, and generate a

R

E

D

histogram of WBC distribution.

N

U

WBC sample RBC sample

Figure 3-5 The dilution process for PD mode

S

E

L

P

I

C

As shown in the figure above, in the PD mode, 20 μL of capillary blood is mixed

N

I

R

P

with the diluent (Out of Machine dilution) to form a diluted sample. The analyzer  M

E

T

aspirates 243 μL of the diluted sample and adds 60 μL to the CRP chamber, and  S

Y

S

then R2 is added after mixing to form a CRP sample.

E

H

T

The remaining 183 μL is added with the diluent to form a secondary diluted

G

N

I

D

sample. This diluted sample will be used in two parts: One part is mixed with the

3

N

A

T

R

diluent to form a secondary diluted sample, which is used for counting RBCs

S

E

R

T

E

P

and platelets and generating a corresponding distribution histogram. The other

A

D

H

N

C

U

part is mixed with the lyse to form another sample, which is used to measure the

2

3

hemoglobin concentration and count WBCs, and generate a histogram of WBC

2

2

0

0

distribution.

CRP PD Mode

Peripheral blood

sample 20μl

A sample with a

diluention ratio of 1:30

WBC/HGB

Measurement

Add Diluent 580μl

Measuring Principles

• WBC counting principle

WBCs are counted and sized by the Coulter method. This method is based on  the measurement of changes in electrical resistance produced by a particle,  which in this case is a blood cell, suspended in a conductive diluent as it  passes through an aperture of known dimensions. An electrode is submerged  in the liquid on both sides of the aperture to create an electrical pathway. As  each particle passes through the aperture, a transitory change in the resistance  between the electrodes is produced. This change produces a measurable  electrical pulse. The number of pulses generated signals the number of particles  that passed through the aperture. The amplitude of each pulse is proportional  to the volume of each particle. Each pulse is amplified and compared to the

The system sucks  183μl sample

Add R1

internal reference voltage channels, which only accepts the pulses of certain  amplitude. If the pulse generated is above the WBC threshold, it is counted as a  WBC.

Diluted Sample

Diluention sample CRP sample

Add R1

Voltage

Detection

Small Holes

Negative Pressure - Constant

Constant

Current

Source

Figure 3-6 The dilution process for CRP PD mode

As shown in the figure above, in the PD mode, 20 μL of capillary blood is mixed

S

Electrodes

Pulse Time

Analvsis Circuits

S

3

R

ET

PA

HC

E

L

with the diluent (Out of Machine dilution) to form a diluted sample. The analyzer

P

I

C

N

I

aspirates 243 μL of the diluted sample and adds 60 μL to the CRP chamber, and

R

P

then R2 is added after mixing to form a CRP sample. The remaining sample is

M

E

T

S

tested for the HCT value to correct the CRP results.

Y

S

E

H

T

G

N

I

D

N

A

T

S

R

E

D

N

U

Figure 3-7 Schematic of counting process

E

L

P

I

C

N

I

R

P

M

E

T

S

Y

S

E

H

T

G

N

I

D

3

N

A

T

R

S

E

R

T

E

P

A

D

H

N

C

U

4 20

5 20

3

R

ET

PA

HC

• Principle of hemoglobin measurement

The hemoglobin concentration is measured by colorimetry. In the WBC

counting chamber, after lyse is added to the diluted sample, the erythrocyte

membrane is dissolved to release hemoglobin, and the latter forms a

hemoglobin complex after being combined with the lyse. On one side of

the WBC counting chamber, the hemoglobin complex solution is illuminated

with an LED monochromatic luminous tube with a center wavelength of 525

nm. On the other side, a photocell receives the transmitted light. The light

intensity signal is first converted into a current signal, then into a voltage

signal and amplified. The hemoglobin concentration (HGB) of the sample (g/

L) is determined by comparing with the voltage generated by the background

intensity of transmitted light measured before the sample is introduced to the

WBC counting chamber (i.e., there is only the diluent in the counting chamber).

This measurement and calculation process is automatically conducted by the

analyzer, and the results will be displayed in the Results area of the “Count”

interface.

WBC Parameters

• WBC count

The analyzer obtains the WBC count (WBC #) (109/L) by directly counting

electric pulses corresponding to WBCs.

WBC= n×109

• WBC grouping

WBCs involve many types of cells that can be divided according to their

volume. The volume of each type of cell varies depending on the diluent and

lyse added, and the hemolysis time. WBCs can be divided into three groups

S

by using certain reagents. They are: lymphocytes, middle size cells (including

E

L

P

I

monocytes, eosinophils, and basophils), and neutrophils.

C

N

I

R

The analyzer calculates the percentage of lymphocytes (Lym%), the percentage

P

M

E

of intermediate cells (Mid%), and the percentage of neutrophils (Gran%)

T

S

Y

according to the histogram of WBC distribution using the following formulae.

S

E

H

T

PL

G

N

Lym%=X 100

I

D

PL + PM + PG

N

A

T

S

R

PM

E

Mid%=X 100

D

N

PL + PM + PG

U

Where, PL is the number of cells in the lymphocyte region, PM is the number  of cells in the intermediate cell region, and PG is the number of cells in the  neutrophilic region. The unit of these three parameters is 109/L.

Based on the percentage of cells in each subpopulation calculated according  to the above formulae, the analyzer automatically calculates the count of  lymphocytes (Lym#), the count of middle size cells (Mid#) and the count of  neutrophils (Gran#) according to the following formulae. The unit of these three  parameters is 109/L.

Lym#= Lym%×WBC

100

Mid#=Mid%×WBC

100

Gran#=Gran%×WBC

100

Where, the unit of Lym%, Mid%, and Gran% is %, and the unit of WBC is 109/L.

• WBC distribution histogram

The analyzer provides WBC volume distribution graphs while providing WBC  count results. The abscissa of the histogram is the volume of WBCs (unit:  fL) and the ordinate is the relative number of WBCs (unit: 109/L). After each  count, the WBC distribution histogram can be viewed in the Result area of  the “Analysis” interface. It can also be viewed in the “Review” interface in a  retrospective manner.

S

E

WBC= n×109/L

L

P

I

C

N

I

R

P

HGB Measurement

M

E

T

S

The analyzer compares the measured voltage with the voltage of the background  Y

S

transmitted light to calculate the hemoglobin concentration (HGB) in g/L.

E

H

T

G

Sample Photocurrent HGB= Constant×Log10 ( )

N

I

Blank Photocurrent

D

3

N

A

T

R

S

E

R

T

E

P

A

D

H

N

C

U

6 20

PG

Gran%=X 100 PL + PM + PG

7 20

RBC/PLT

Measurement

Electrical Impedance Method

RBCs/PLTs are counted and sized by the Electrical Impedance method. This  method is based on the measurement of changes in electrical resistance  produced by a particle, which in this case is a blood cell, suspended in a  conductive diluent as it passes through an aperture of known dimensions. A  pair of electrodes is submerged in the liquid on both sides of the aperture to  create an electrical pathway. As each particle passes through the aperture, a  transitory change in the resistance between the electrodes is produced. This  change produces a measurable electrical pulse. The number of pulses generated  represents the number of particles that passed through the aperture. The  amplitude of each pulse is proportional to the volume of each particle.

Diluted sample

Negative Pressure

Aperture

• Hematocrit (HCT), mean corpuscular hemoglobin content (MCH), mean   corpuscular hemoglobin concentration (MCHC)

HCT in %, MCH in pg, and MCHC in g/L can be calculated using the following  formulae.

HCT=RBC × MCV

10

MCH=HGB

RBC

MCHC=HGB

HCTX 100

Where, RBC count is in 1012/L, MCV is in fL, and HGB is in g/L. • RBC distribution width - coefficient of variation (RDW-CV)

Voltage

Impulse Time

Electrode

Circuits

Constant Current Source

RDW-CV is derived from the distribution histogram of RBCs. It is the coefficient  of variation of the volume distribution expressed as a percentage.

• RBC distribution width - standard deviation (RDW-SD)

RDW-SD is the width of the histogram at the 20% peak of the histogram of the  distribution of RBCs in fL, as shown in Figure 3-5.

Figure 3-8 Electrical Impedance Method

100%

20%

Derivation of RBC-Related Parameters

S

S

RDW-SD

E

E

• RBC

L

L

P

P

I

I

3

R

ET

PA

HC

C

The analyzer counts RBCs (RBC #) in 1012/L by directly counting electrical

N

I

R

pulses corresponding to RBCs.

P

M

E

T

S

RBC= n×1012/L

Y

S

E

H

T

G

• Mean corpuscular volume (MCV)

N

I

D

Based on the RBC distribution histogram, the MCV can be calculated in fL.

N

A

T

S

R

E

D

N

U

Figure 3-9 Schematic Diagram

C

N

I

R

P

• Histogram of RBC distribution

M

E

T

The analyzer provides the RBC volume distribution graph while giving the

S

Y

S

RBC count results. The graph that can represent the distribution of the cell

E

H

T

population is called the RBC distribution histogram. The abscissa of the

G

N

I

histogram is the RBC volume (unit: fL) and the ordinate is the relative number

D

3

N

of RBCs (unit: 1012/L) . After each count, you can view the RBC distribution

A

T

R

S

E

R

T

histogram in the "Results" area of the "Analysis" interface, or you can enter the

E

P

A

D

"Review" interface to view the RBC distribution histogram in a retrospective

H

N

C

U

8 20

manner.

9 20

3

R

ET

PA

HC

Derivation of PLT-Related Parameters

• PLT

The analyzer counts platelets (PLT) in 109/L by directly counting electrical pulses

corresponding to platelets.

PLT= n×109/L

• Mean platelet volume (MPV)

Based on the histogram of platelet distribution, MPV is calculated in fL.

• Platelet distribution width(PDW)

Assuming that the peak height is 100%, the distribution width at the 20%

frequency level is PDW in fL.

• Plateletcrit (PCT)

The analyzer calculates PCT in % using the following formula. Where, PLT is in

109/L and MPV in fL.

PCT=PLT × MPV

10000

• Platelet-large cell ratio (P-LCR)

According to the histogram of platelet distribution, P-LCR is calculated in %.

• Platelet-large cell count (P-LCC)

Based on P-LCR and the platelet count, P-LCC is calculated in 109/L.

P-LCC= PLT x P - LCR

S

E

L

• Histogram of platelet distribution

P

I

C

The analyzer provides the platelet volume distribution graph while giving

N

I

R

P

the platelet count results. This graph that shows the distribution of this cell

M

E

subpopulation is called the platelet distribution histogram. The abscissa of the

T

S

Y

histogram is the platelet volume (unit: fl) and the ordinate is the relative number

S

E

of platelets (unit: 109/ L). After each count, you can view the platelet distribution

H

T

G

histogram in the "Results" area of the "Analysis" interface, or you can enter the

N

I

D

"Review" interface to view the platelet distribution histogram in a retrospective

N

A

T

manner.

S

R

E

D

N

U

CRP

Measurement Rinse

The analyzer measures the concentration of CRP by immunoturbidimetry. The  HCT correction technique is used to obtain the CRP concentration with higher  accuracy.

Immunodiffusion Turbidimetry Method

The analyzer emits monochromatic light at a certain wavelength into the CRP  chamber, and the light scatters when encountering the antigen (CRP)-antibody  complexes. The analyzer receives the scattered light through a photocell and  then amplifies the light intensity signal and converts it into a voltage signal.  The analyzer obtains the CRP concentration of the sample by comparing the  measured voltage signal with the voltage generated by the background light  intensity measured before adding the sample to the CRP chamber. The intensity  of the scattered light is directly proportional to the content of the complexes,  that is, the more antigen (CRP) in the sample, the more complexes form and the  stronger the scattered light.

HCT Corrected CRP Results

The analyzer uses the HCT correction technique to calculate the CRP

concentration in serum, thus improving the accuracy of the results.

C-reactive Protein Parameters

C-reactive protein (CRP): First, a standard curve is obtained by measuring the  antigen calibrators, and then the blood samples are measured. Second, the  CRP value before HCT correction is obtained from the standard curve, and then  the concentration value of CRP in mg/L is calculated according to the following  formula.

CRP value on the standard curve

1- HCT CRP=

S

E

Hypersensitive C-reactive protein (Hs-CRP): Concentration value of Hs-CRP. Unit:

L

P

I

mg/L.

C

N

I

R

P

M

E

T

S

Y

S

E

During each counting process, the analyzer automatically flushes the

H

T

components through which the sample flows, ensuring that there is no sample

G

N

I

D

residue in the fluidic components.

3

N

A

T

R

S

E

R

T

E

P

A

D

H

N

C

U

0 30

1 30

4

R

ET

PA

HC

Introduction

Installation

Requirements

R

E

Z

Y

L

A

N

A

R

U

O

Y

G

N

I

L

L

A

T

S

N

I

CHAPTER 4

INSTALLING

YOUR ANALYZER

Personnel who are not authorized or trained by Zybio may cause personal injury  or analyzer damage when unpacking or carrying out the installation process. Do  not unpack or install the analyzer in the absence of the personnel authorized by  Zybio.

The analyzer has been strictly tested before shipment. To avoid collision during  transportation, the analyzer has been carefully packed before transportation.  When the analyzer arrives, please carefully check the carton to see if there is any  physical damage. If there is any damage, please immediately notify the after-sales  service department of Zybio or the local agent.

Before installation, the operator must ensure that the following requirements for  the space, power supply, environment and fuses are met.

Space Requirements

Ensure there is sufficient space for maintenance and repair. Considering the heat  dissipation of the instrument and the non-extrusion of the fluidic components  behind the analyzer (for the normal flow of reagents), the following requirements  shall be met:

• a space of ≥ 30 cm left between the left and right doors of the analyzer and the  walls;

• a space of ≥ 10 cm left between the rear panel of the analyzer and the wall; • The installing table (or floor) can bear a weight of at least 40 Kg. • Make sure there is enough room on the work table surface and below the  analyzer to place the reagents, such as diluents, and waste buckets.

Power Requirements

POWER VOLTAGE POWER FREQUENCY INPUT POWER FUSE

MAIN UNIT

100V - 240 V (50/60 Hz)± 1Hz ≤200 VA 250 V T6.3 AH

• The analyzer must be used under good grounding conditions.

• The operator must use a fuse of the specified specifications.

• Verify that the input voltage meets the instrument requirements.

• The use of a power strip may introduce additional electrical interference and  result in erroneous analysis results. Please place the analyzer near the power  outlet to avoid using a power strip.

• Please use the supplied power cord. Using other power cords can damage the  analyzer or cause erroneous analysis results.

Environmental Requirements

ENVIRONMENTAL REQUIREMENTS STORAGE OPERATION

Ambient temperature -10℃-40℃ 15℃-35℃

Relative humidity 10%-90% 20%-85%

Atmospheric pressure 50kPa-106kPa 70kPa-106kPa

• The environment should be free from dust, mechanical vibration, major noise  sources and power interference.

• It is recommended that the electromagnetic environment of the laboratory be  evaluated before running the equipment.

• Please use a dedicated power outlet. Do not use the same power outlet as  air-conditioners, refrigerators, ultrasound systems, etc. that are likely to emit  interference signals.

• Do not place the device near strong electromagnetic interference sources so as  not to affect the normal operation of the device.

R

E

• Do not place the device near brush-type motors, flashing fluorescent lights, and

Z

Y

L

electrical contact devices that are frequently switched on/off.

A

N

A

R

• Avoid direct sunlight or heat and wind sources.

U

O

Y

4

• Choose a well-ventilated location.

G

N

R

I

L

E

L

T

• Maintain a good grounding environment.

A

P

T

A

S

H

N

C

I

2 30

• Indoor use only.

3 30

4

R

ET

PA

HC

• The analyzer may not be used in presence of flammable substance and/or

explosives

• If the room temperature exceeds the normal operating temperature range of

the analyzer, the instrument temperature may exceed the limit and the analytical

results obtained will be unreliable.

Handling

• Personnel who are not authorized or trained by Zybio may cause personal injury or

damage to the main unit when unpacking or carrying out the installation process. Do

not unpack or install the main unit in absence of the authorized personnel of Zybio.

• During transportation, in order to avoid damage to the sampling component,

the moving components are immobilized with clips/tying tapes when the

instrument leaves the factory. The clips/tying tapes must be removed before

using the instrument.

The analyzer shall be transported and installed by the personnel authorized by

Zybio. Do not move or install the analyzer without contacting the after-sales

service department of Zybio or the local agent.

R

E

Z

Y

L

A

N

A

R

U

O

Y

G

N

I

L

L

A

T

S

N

I

Connecting the  Analyzer System

Make electrical and reagent connections as shown in the figure below. The  operator must verify that the connections are in place and secure.

1

2

Figure 4-1 Electrical connection diagram

1-Network port 2-Power component

• The operator is obliged to comply with the relevant national and regional  regulations regarding the discharge and processing of expired reagents, waste  liquids, waste samples, consumables, etc.

• Reagents may irritate the eyes, skin and mucous membranes. When the

operator handles reagent-related articles in the laboratory, he/she shall comply  R

with the laboratory safety practices and wear personal protective equipment

E

Z

Y

L

(such as laboratory protective clothing, gloves, masks, etc.).

A

N

A

• Once the reagent contacts the skin, rinse with plenty of water immediately. If

R

U

O

necessary, please seek medical treatment. Once the reagent contacts the eyes,

Y

4

G

immediately rinse with plenty of water and seek medical treatment.

N

R

I

L

E

L

T

A

P

T

A

S

H

N

C

I

4 30

5 30

1

2

3

Figure 4-2 Diagram of external reagent connections

2

3

Figure 4-3 Diagram of internal reagent connections

1-Connection to diluent

2/3-Connection to a waste bucket

1-Lyse 2-R2 kit

3-CRP measuring cup

4

R

ET

PA

HC

R

EZ

YL

A

N

A

RU

O

Y

G

NI

LL

AT

S

NI

• Please ensure that the length of the diluent and waste conduits does not

exceed 1500mm.

• The top height of the waste and diluent buckets should be lower than the table

top on which the instrument is placed.

R

E

Z

Y

L

A

N

1

A

R

U

O

Y

4

G

N

R

I

L

E

L

T

A

P

T

A

S

H

N

C

I

6 30

7 30

4

R

ET

PA

HC

Notes

Installing the

thermal paper • Remove the protective paper from the recorder before installing the thermal

paper for the first time.

The thermal paper is installed as follows:

1. Open the recorder door outward.

2. Load the thermal paper into the paper chamber in the direction shown below, with

the paper’s heading end outside the paper outlet.

3. Close the recorder door.

4. Check the position of the thermal paper to ensure that the thermal paper is

aligned with the paper outlet.

Figure 4-4 Installing the thermal paper

• Select qualified 50mm thermal paper.

• During the printing process of the recorder, the thermal paper cannot be pulled

outwards by force; otherwise the recorder may be damaged.

• Do not leave the recorder door open unless you are changing the paper or

troubleshooting the recorder.

• Thermal paper installation errors may cause paper jams or printing failure.

R

E

Z

Y

L

A

N

A

R

U

O

Y

G

N

I

L

L

A

T

S

N

I

• The reagents designated by Zybio should be used, otherwise the test results  will be unreliable and operation may cause damage to the instrument.

• Attention should be paid to the expiry date of the reagents. Expired reagents  may not be used. The use of expired reagents will lead to unreliable test results.

• After the reagent is connected with the analyzer, the reagent bottle cap must be  replaced to prevent the reagent from being polluted.

• The analyzer performance may be undermined if it has been placed in  environment of high dustiness.

• The surface of the analyzer shall be cleaned and sterilized regularly with alcohol  (75%).

• The probe wipe block of the analyzer (see Figure 4-5 Front of the analyzer) shall  be wiped with alcohol (75%) regularly.

• Blood collection and sample preparation shall be carried out in accordance  with the specified methods. Inappropriate blood collection procedures may cause  harm.

• If any hoses or parts filled with liquid are aged or worn during use, please stop  using them immediately and contact the user’s service personnel for inspection or  replacement in a timely manner.

• During the use of the instrument, care should be taken not to compress with  heavy objects or bend the connection tubing of the reagents (including the  diluent, lyse and waste liquid).

R

E

Z

Y

L

A

N

A

R

U

O

Y

4

G

N

R

I

L

E

L

T

A

P

T

A

S

H

N

C

I

8 30

9 30

5

R

ET

PA

HC

Introduction

R

E

Z

Y

L

A

N

A

R

U

O

Y

G

N

I

T

A

R

E

P

O

CHAPTER 5

OPERATING

YOUR ANALYZER

This chapter introduces the routine operation process from the start-up to the  shut-down of the analyzer, detailing the sample analysis process in different  working modes.

The routine operation process is as follows:

Preparation

Start-up

Daily Quality Control

Sample Preparation

Preparation

Sample Analysis

Report and Management

Initial Checks

Before powering on the main unit, the operator must check the following to  ensure that the system is ready.

All articles (samples, controls, calibrators, reagents, waste liquid, etc.) and the  areas that come into contact with these substances pose potential biological  risks. When the operator comes into contact with such articles and areas in  the laboratory, you shall comply with the laboratory safety practices and wear  personal protective equipment (such as laboratory protective clothing, gloves,  masks, etc.).

• The operator is obligated to comply with the relevant national and regional  regulations regarding the discharge and processing of reagents, waste liquid,  waste samples, consumables, etc.

• Reagents may irritate the eyes, skin and mucous membranes. When the  operator handles reagent-related articles in the laboratory, he/she shall comply  with the laboratory safety practices and wear personal protective equipment  (such as laboratory protective clothing, gloves, masks, etc.).

• Once the reagent contacts the skin, rinse with plenty of water immediately. If  necessary, please seek medical treatment.

• Keep your clothing, hair, and hands at a certain distance from the moving parts

• The operator shall use the reagents specified by Zybio and store and use them  in strict accordance with their MANUAL.

• Before using the analyzer, verify that the reagents are properly connected.

• The reagents should be allowed to stand motionless for a period of time until

R

they become stable.

E

Z

Y

L

A

N

A

R

U

O

Y

5

G

N

R

I

E

T

T

A

P

R

A

E

P

H

C

O

0 40

Shutdown

1 40

5

R

ET

PA

HC

Startup and Login

R

E

Z

Y

L

A

N

A

R

U

O

Y

G

N

I

T

A

R

E

P

O

• Check the waste bucket

The operator must put a waste bucket in place and ensure that it is empty   before starting the machine each day.

• Check the fluidic components and power supply

Check the reagent and waste tubing for bending or insecure connections.  Check that the power plug of the main unit is firmly inserted into the power socket.

• Check the recorder and printer (optional)

Check the recorder and printer for insufficient paper or inappropriate installation.

• Start up the main unit:

1. Switch the power I/O switch on the back of the analyzer to position "I" and the   power indicator will illuminate.

2. Verify that the indicator on the main unit is On.

3. In the login dialog box, enter the current user's username and password in the   "User name" and "Password" boxes.1

Figure 5-1 Login dialog box

4. The analyzer performs self-check and power-on initialization in sequence. The   time required for the analyzer to initialize the fluidic components varies

Daily Quality  Contro

Sample

Preparation

• If analysis is performed when the analyzer reports “Background Abnormal”, the  analyzer will yield unreliable results. Please handle this error according to Chapter  11 Troubleshooting Your Analyzer.

• The system judges the operator's privileges as Administrator or Common User  according to the user name and password used for login, and then enables  different functions in each interface according to the user's privilege.

• To switch users, click “Logout" in the menu, enter the user name and password  in the login dialog box, and click “Login" to log into the software interface as a  new user.

1. The initial username and password of the Administrator default to Admin.  2. If the software fails to run after several consecutive attempts, please contact   the after-sales service department of Zybio or your local agent.

3. Please verify that the date/time of the device is valid after startup.

Before carrying out sample analysis, QC analysis shall be carried out on the  analyzer daily to ensure that the analyzer obtains reliable analysis results. Refer to  Chapter 7 Quality control for specific QC analysis methods.

The samples measured by the instrument are: whole blood samples, and  pre-diluted samples

• Samples should be prepared according to the procedures recommended by  the reagent manufacturer.

• All kinds of samples must be thoroughly mixed.

R

E

Z

Y

L

All articles (samples, controls, calibrators, reagents, waste liquid, etc.) and the  A

N

areas that come into contact with these substances pose potential biological

A

R

U

risks. When the operator comes into contact with such articles and areas in the  O

Y

5

laboratory, he/she shall comply with the laboratory safety practices and wear

G

N

R

I

personal protective equipment (such as laboratory protective clothing, gloves,

E

T

T

A

P

R

etc.).

A

E

P

H

C

O

according to the previous shutdown conditions.

2

3

4

4

0

0

5

R

ET

PA

HC

• Do not directly come into contact with the patient's blood samples.

• Do not reuse disposable supplies.

• The operator should use clean EDTAK2 anticoaguled vacuum blood collection

tubes, silicified glass/plastic test tubes, centrifuge tubes and borosilicate glass

capillary tubes.

• Disposable supplies such as vacuum blood collection tubes, centrifuge tubes

and capillary tubes used in blood collection must be in accordance with the

specifications specified by the manufacturer.

Whole Blood Samples

1. Venous blood samples are collected using EDTAK2anticoagulated vacuum tubes.

1. Quickly mix venous blood in the tube with the anticoagulant thoroughly.

• To ensure the accuracy of the results, the sample volume of capillary WB

should not be less than 120 μL.

• Samples for WBC differential counting or platelet counting should be stored at

room temperature and analyzed within 8 hours.

• If the sample is kept in a refrigerator at 2℃- 8℃, it can be analyzed within 24

hours. Refrigerated samples should be left at room temperature for at least 30

R

minutes before analysis.

E

Z

Y

L

A

• Samples placed for a certain period of time need to be remixed before analysis.

N

A

R

U

• Please complete the analysis 5 minutes to 2 hours after sample collection.

O

Y

G

N

I

T

A

R

E

P

O

Sample Analysis

Pre-diluted Samples

1. Click on the mode switch icon to change the analysis mode from “WB” to "PD". 2. Click the "Diluent" icon to display the diluent dispensing prompt dialog box.

Place a clean centrmanualge tube under the sample probe and press the Aspirate  Key to start adding the diluent (580 μL). While adding the diluent, the prompt box  displays “Adding diluent …" and a progress bar appears.

Collect 20 μL venous or capillary blood and quickly inject it into a centrmanualge  tube filled with the diluent. Replace the cap and mix thoroughly. After the PD  sample is prepared, click the “Cancel" button to exit diluent dispensing.

• The operator can also use a pipette to draw 580 μL of diluent.

• The prepared diluent should be kept away from dust and volatilization  prevented, otherwise analysis errors will occur.

• After the capillary blood reacts fully with the diluent, it needs to be left for 3  minutes and remixed before analysis.

• It is recommended that the analysis be completed within 30 minutes of sample  dilution.

• Samples unused for a certain period of time need to be remixed before  analysis.

• Each laboratory shall evaluate the stability of the sample analysis results in the  PD mode according to its own sample number, sample collection method and  technical level.

Click the "Analysis" button to enter the "Analysis" interface. Click the “Mode  switch" button in this interface to select the "WB" or “PD" mode.

Enter the sample information

R

E

The analyzer provides two methods to enter the sample information: Sample ID

Z

Y

L

entry and all information entry.

A

N

A

R

If the operator wishes to enter the sample information after analysis, he/she can  U

O

Y

skip the introduction in this section and enter the sample information according to

5

G

N

the Sample ID and the result saving time when reviewing the sample results. See

R

I

E

T

T

A

P

Chapter 6 Review of results for the method.

R

A

E

P

H

C

O

Once the sample information entry method has been set in the "Setup → Auxiliary"

4

5

4

4

interface according to Chapter 9 Setup, the sample information can be entered in

0

0

the "Analysis" interface.

5

R

ET

PA

HC

Entering all information

When the entry method of the next sample is set to "All Information", click

“Analysis" and “Next Sample" to open the all information entry dialog box, as

shown in the following figure. The operator can enter the complete sample

information for the next sample in the dialog box.

Figure 5-2 All information entry dialog box

• Sample mode selection

Click "WB"/"PD" to select the sample mode.

• Measurement mode

Click "CBC"/"CRP"/"CBC + CRP" mode to select the measurement mode.

• Enter Sample ID

Enter the Sample ID in the "ID" box.

R

• Enter patient name

E

Z

Y

Enter the patient's name in the "Name" box.

L

A

N

A

• Select patient gender

R

U

O

Select the gender of the patient from the "Gender" drop-down list. There are

Y

G

three options: “Male", “Female" and “Unknown". The default option is “Unknown".

N

I

T

A

R

E

P

O

• Enter patient age

The analyzer provides five time units for various age groups: by “Years", by  “Months", by “Weeks", and by “Days" and by “Hours". They are applicable,  respectively, to: people aged over one year, people aged a full month but under  two years, people aged a full week but under ten weeks, people aged under a  full month and people aged under 48 hours. The operator can select the time  unit of the patient's age accordingly.

In the “Age" drop-down list, select the time unit of the age in “Years" , “Months" ,  “Weeks" , “Days" or “Hours" and enter the patient's age in the entry box in front  of the time unit.

• Enter patient age

The analyzer provides five time units for various age groups: by “Years", by   “Months", by “Weeks", and by “Days" and by “Hours". They are applicable,   respectively, to: people aged over one year, people aged a full month but under   two years, people aged a full week but under ten weeks, people aged under a   full month and people aged under 48 hours. The operator can select the time   unit of the patient's age accordingly.

• In the “Age" drop-down list, select the time unit of the age in “Years" , “Months" ,   “Weeks" , “Days" or “Hours" and enter the patient's age in the entry box in front   of the time unit.

R

E

Z

Y

L

A

N

A

R

U

O

Y

5

G

N

R

I

E

T

T

A

P

R

A

E

P

H

C

O

6 40

7 40

5

R

ET

PA

HC

• Enter the birth date

Enter the patient's birth date in the "Birth date" box. The date format is

consistent with the system date format.

• Enter the deliverer

Enter the name in the "Deliverer" box or select the name in the "Deliverer" drop-

down list (when there is a record in the drop-down list).

• Enter the delivery time

Enter the draw time in the "Delivery Time" box.

• Enter the patient ID

Enter the patient ID in the "Patient ID" box.

• Select the patient type

Select the patient's type from the "Patient Type" drop-down list. There are four

options: Outpatient, Inpatient, Checkup, and Emergency.

• Enter the department name

Enter the department name in the "Dept." box, or select the department name

in the "Dept." drop-down list (when there is a record in the drop-down list).

• Enter the bed number

Enter the patient’s bed number in the "Bed No." box.

• Enter the draw time

Enter the draw time in the "Draw Time" box.

• Enter remarks

Enter necessary remarks in the "Remarks" box.

• OK

After entering the sample information, click OK to save the entry and return to

the "Analysis" interface.

• Cancel

After entering the sample information, click “Cancel" to return to the "Analysis"

R

interface and discard the entry.

E

Z

Y

L

• Enter Sample ID

A

N

A

When the entry method of the next sample is set to Sample ID only, click "Next

R

U

Sample" in the “Analysis" interface to open the ID entry dialog box.

O

Y

G

N

I

T

A

R

E

P

O

Edit current sample information

Click on the sample information area in the "Analysis" interface and the “Graph  Review” interface to open the sample information editing dialog box to edit  the information of the current sample. Sample information for background and  validated samples cannot be edited.

Sample Analysis Steps

All articles (samples, controls, calibrators, reagents, waste liquid, etc.) and the  areas that come into contact with these substances pose potential biological  risks. When the operator comes into contact with relevant articles and areas in  the laboratory, he/she shall comply with the laboratory safety practices and wear  personal protective equipment (such as laboratory protective clothing, gloves,  etc.).

R

E

Z

Y

L

A

N

• The sample probe is sharp, and it may carry blood samples, controls and

A

R

U

calibrators that are potentially biologically hazardous. Therefore, the operator shall  O

Y

not touch the sample probe.

5

G

N

R

I

E

T

T

A

P

R

A

E

P

H

C

O

8 40

9 40

• Do not reuse disposable supplies.

• When the sample probe is aspirating, it needs to be fully immersed in the  sample and its tip should be kept at a certain distance from the bottom of  the container, otherwise the sample volume aspirated may be insufficient and  inaccurate.

• The operator should avoid blood splashing caused by the contact between the  test tube wall and the sample probe.

Report Management

• Save the analysis results

The analyzer automatically saves the results. When the number of sample   results reaches its upper limit, the newly obtained results will automatically   overwrite the oldest results.

Text alarms

Flags Alarm information Description

Leucopenia WBC count significantly low

Leucocytosis WBC count significantly high

Granulopenia Granulocyte count significantly low

Granulocytosis Granulocyte count significantly high

Sample analysis

WB samples are analyzed according to the following steps:

WBC Flag

WBC Abnormal

There may be nucleated RBC, abnormal  lymphocytes, immature cells, primitive cells  or other abnormalities

1. Verify that the Analysis status in the system status area is ready and the   working mode is “WB” or “PD”.

2. Place the prepared WB sample under the sample probe so that the sample   probe can aspirate the mixed sample.

Lymphopenia Lymphocyte count significantly low Lymphocytosis Lymphocyte count significantly high Increased Intermediate Cells Intermediate cell count significantly high There may be small RBCs, large RBCs,

3. Press the Aspirate Key to start the sample analysis process. At this point, the   blue flashing status of the analyzer indicator indicates that the sample analysis   is in progress.

RBC Abnormal

Hemoglobin Abnormal/ Interference?

anisocytosis, RBC agglutination, double  peaks on the histogram and other  abnormalities

There may be abnormal hemoglobin, RBC  agglutination, etc.

RBC/HGB Flag

4. The sample probe automatically sucks in the sample and then lifts itself up,

while buzzing. After the sample probe is lifted, the operator can remove the

sample. Then, the sample probe adds the aspirated sample to the counting

chamber. The analyzer automatically performs sample analysis.

5. After the analysis is finished, the sample probe is reset and ready for the next

Microcytosis RBC volume is low Macrocytosis RBC volume is high Anemia Anemia

Erythrocytosis RBC count significantly high  There may be small RBCs, RBC

sample analysis. The results will be displayed in the results area of the interface.

PLT Flag

Platelets Abnormal

fragments, giant platelets, platelet  aggregation and other abnormalities

Simultaneously, the number of the next sample is automatically increased by one.

R

E

6. If automatic printing is set to On, the analyzer will automatically print the analysis

Z

Y

L

A

report as configured. If auto communication is set to On, the analyzer will

N

A

automatically upload the sample analysis results and sample and patient

R

U

O

Thrombopenia Platelet count significantly low

R

E

Thrombocytosis Platelet count significantly high

Z

Y

L

A

N

A

Validate

R

U

O

5

R

ET

PA

HC

information that meet the communication conditions to the LIS system.

Y

G

N

I

7. The remaining samples are analyzed by following the same procedure.

T

A

R

E

P

O

Validate the results of the current sample.

Y

5

G

N

R

I

E

T

T

A

P

R

A

E

P

H

C

O

0 50

1 50

5

R

ET

PA

HC

Sleep

Shutdown

R

E

Z

Y

L

A

N

A

R

U

O

Y

G

N

I

T

A

R

E

P

O

Once the fluidic-related operations have stopped for a time period sufficient  to trigger the sleep mode as set by the operator in the setting interface, the  analyzer enters the sleep state.

After the main unit enters sleep mode, the lower-left corner of the interface  shows “The analyzer is sleeping, please click Aspirate Key to wake up”.

Perform the shutdown procedure before powering off the analyzer each day,  which includes the following steps:

All articles (samples, controls, calibrators, reagents, waste liquid, etc.) and the  areas that come into contact with these substances pose potential biological  risks. When the operator comes into contact with relevant articles and areas in  the laboratory, he/she shall comply with the laboratory safety practices and wear  personal protective equipment (such as laboratory protective clothing, gloves,  etc.).

• The sample probe is sharp, and it may carry blood samples, controls and  calibrators that are potentially biologically hazardous. Therefore, the operator  shall not touch the sample probe.

• To ensure the stability of the analyzer and the accuracy of the results, please  shut down the analyzer as required after 24 hours of continuous operation.

• The operator must implement the required shutdown procedure to shut down  the machine according to the following steps.

• Please do not forcibly turn off the power supply during shutdown.

• If there is a failure that affects shutdown, the analyzer will return to the state  before shutdown and give an alarm. See Chapter 11 Troubleshooting for the  workaround.

1. Click the “Shutdown" button at the bottom-left of the interface;

2. Select “Yes”, place the probe cleanser under the sample probe, and press the   Aspirate Key. The analyzer automatically performs probe soaking.

3. After the analyzer automatically performs the shutdown process, it prompts   "Please turn off the power!”. Then, turn off the analyzer’s power switch.

4. Empty the waste bucket and dispose of the waste properly.

• The operator is obliged to comply with the relevant national and regional  regulations regarding the discharge and processing of expired reagents, waste  liquids, waste samples, consumables, etc.

R

E

Z

Y

L

A

N

A

R

U

O

Y

5

G

N

R

I

E

T

T

A

P

R

A

E

P

H

C

O

2 50

3 50

6

R

ET

PA

HC

Introduction

Sample review

S

T

L

U

S

E

R

E

L

P

M

A

S

G

N

I

W

E

I

V

E

R

CHAPTER 6

REVIEWING

SAMPLE RESULTS

After each sample analysis, the analyzer automatically stores the results in  the sample library. The sample library can store up to 40,000 results including  parameter results and histograms.

The operator can review all of the sample parameter results and histograms  stored in the sample library and search library by listing or by single samples with  histograms.

• Effective backups should be performed to prevent data loss in the event of  hardware or software failure.

Click “Review" in the menu to review the analyzed record. The serial number,  Sample ID, sample status, and analysis parameters are displayed in sequence in  the sample results area as a list.

Search

Click the "Search" button to open the dialog box shown below.

Define the search criteria by entering content in the corresponding edit boxes or  selecting from the drop-down list.

Click “OK" to close the dialog box and start the search. The search results will be  displayed in the list area.

Graph

The operator can click the "Graph Review" button in the “Sample review"  interface to browse the detailed analysis results of each sample.

S

T

L

U

S

E

R

E

L

P

M

A

S

6

G

N

R

I

E

W

T

E

P

I

A

V

H

E

C

R

4 50

5 50

Edit result

The operator can click on a result entry in the "Graph" interface and click the  "Edit Result" button to open the interface shown in the following figure:

Trend graph

Click the “Trend Graph" button to see the trend graph of sample results.

From the trend graph, you can view the sample CV value over a period of time.

Modify some of the results of this sample, and click the "OK" button to save.

Then, return to the "Graph Review" interface. The parameter results on the

6

R

ET

PA

HC

interface will be automatically recalculated and refreshed based on the modified

results.

Edit information

In the "Review" interface, select an entry, click the “Graphic review" button,

and then click the sample information area to open the interface shown in the

following figure:

S

T

L

U

S

E

R

E

L

P

M

A

S

G

N

I

W

E

I

V

E

R

Validate/Cancel Validation (for administrators only)

• Validate sample data

After selecting one or more unvalidated sample records in the "Review"   interface, click the "Validate" button, and the word “Validated" will appear in the   sample status bar of the sample records.

S

T

L

U

S

E

R

E

L

P

M

A

S

6

G

N

R

I

E

W

T

E

P

I

A

V

H

E

C

R

6 50

7 50

6

R

ET

PA

HC

• Cancel validation

After selecting one or more validated sample records in the "Review" interface,

click the "Cancel" button, and the word “Validated" in the sample status

disappears.

Print

Select the sample records to be printed in the list area, and then click the “Print"

button to print. For the samples already printed, the word “Printed" will appear in

the sample status bar of the "Review" interface.

LIS

1. Click the "LIS " button in the “Review” interface.

2. Select the "Check Record" radio button.

3. Click "OK" to close the dialog box and start communicating. The selected results

can be pushed to the data management software.

Export

1. Insert a USB flash disk into the USB port on the back of the instrument.

2. Click the "Export" button to bring up a dialog box.

3. In the "Export Range" area, select “Selected Records” or “All Records”.

Delete

1. Select the sample records to be deleted in the list area.

2. Click the "Delete" button.

3. Click "OK" to delete the selected sample records and close the dialog box.

S

S

T

T

L

L

U

U

S

S

E

E

R

R

E

E

L

L

P

P

M

M

A

A

S

S

6

G

G

N

N

R

I

I

E

W

W

T

E

E

P

I

I

A

V

V

H

E

E

R

C

R

8 50

9 50

7

R

ET

PA

HC

Introduction

S

M

A

R

G

O

R

P

C

Q

E

H

T

G

N

I

S

U

CHAPTER 7

USING THE

QC PROGRAMS

Hematology analyzers may produce somewhat erroneous results after a long  period of use. The presence of errors can lead to wrong or unreliable analysis  results. The QC program provides an effective method for detecting possible

errors. The operator can only effectively eliminate the influence of errors on  the results if he/she is familiar with the theory of QC and masters the practical  operation methods.

To ensure the reliability of the sample analysis results, it is recommended that  the operator conduct QC on the analyzer with low, medium, and high levels  of controls each day. When a new lot of controls is to be used, the new lot of  controls and the existing controls are used in parallel for 5 days, two runs a day.  The results should fall in the reference range specified in the MANUAL of the  controls.

The analyzer provides two QC methods. Click on the “QC” menu and select “L-J  QC” or “X-B QC”.

• The operator should use the specified controls and reagents, and store and use  them strictly in accordance with their MANUALs.

L-J QC

QC Settings

Before the analysis of a new lot of controls, a QC file needs to be set up for each  lot of controls.

1. Click "Quality Management" > "L-J QC" > "Setting".

2. Enter the QC setting interface shown below.

Enter QC Information

1. Enter the L-J QC setting interface.

2. Click the "Setting" > "New" buttons, or select a QC file with no QC count   results and click the “Modify" button.

3. Manually enter the Lot.

• The Lot cannot be blank. The entry should be 1-16 characters, and special  S

M

characters, numbers and letters are allowed, but Chinese characters are not

A

R

G

supported.

O

R

P

C

Q

7

E

H

R

T

E

T

G

P

N

A

I

H

S

C

U

0 60

1 60

7

R

ET

PA

HC

4. Select the level of the controls.

5. Enter the expiry date of the control lot.

6. Select the appropriate “QC Type” from the drop-down list.

7. Select the QC mode for analyzing the controls.

8. Select the measurement mode for analyzing the controls.

9. Setting QC ID: If the operator is accustomed to placing the controls into daily

samples for analysis, a special number can be set here for the controls. If the

instrument recognizes this special number during the analysis of daily samples,

it will automatically recognize it as a control. After the analysis, the test results

will be stored in the QC file corresponding to this number.

10. According to the target value table of the corresponding lot, enter the

reference value and limits respectively in the edit boxes after the parameter

subject to QC.

S

11. Click “Save” to save the entered QC information.

M

A

R

G

O

R

P

C

Q

E

H

T

G

N

I

S

U

Set Limits

If you want to adjust the display of the limits, you can follow the following steps: 1. Click the "Set Limits" button.

2. If you want the limits to be displayed as an absolute value, click "By SD (#)"; if   you want the limits to be displayed as a percentage, click "By CV (%)". 3. Click the "OK" button to save the settings.

QC Count

The operator can choose one of the following two methods for QC analysis  according to his/her preference:

• Use controls to perform QC analysis in the QC counting interface

• Place controls in daily samples and perform QC analysis in the sample counting  interface

S

Use controls to perform QC analysis in the QC counting interface

M

A

After QC editing, you can select one of the following methods for QC analysis

R

G

O

according to the selected QC mode:

R

P

• WB

C

Q

7

• PD

E

H

R

T

E

T

G

P

N

A

I

H

S

C

U

2 60

3 60

7

R

ET

PA

HC

• Running a QC in the event of a error may result in incorrect analysis results. If

a error alarm occurs during QC analysis, be sure to perform QC analysis after

troubleshooting.

• Sample agglutination may result in inaccurate analysis results. Before

the analysis, please check the controls for agglutination. If there is sample

agglutination, please handle according to the relevant operating requirements of

the laboratory.

1. Click "QC" > "L-J QC" to enter the QC counting interface.

• Verify that the level of the control to be analyzed is as shown in the selected

empty file and that the control to be analyzed has not expired.

• The expiry date field of expired controls is indicated in red.

2. P repare controls in accordance with the MANUAL of the controls.

S

M

A

3. Perform QC analysis:

R

G

O

1) Verify that the QC mode is "WB" or "PD" and the main unit indicator is blue.

R

P

2) Mix and manipulate the controls according to their MANUAL and thoroughly

C

Q

E

mix the samples.

H

T

3) Place the control object under the sample probe and click on the Aspirate

G

N

I

Key to start counting.

S

U

4. After the analysis, the QC results are automatically saved to the QC file, and  the newest QC results are displayed in the current interface.

• Each QC file stores up to 200 QC results.

5. If necessary, repeat the above steps to continue the QC analysis.

Place controls in daily samples and perform QC analysis in the sample  counting interface

After setting a special "QC ID" for the control in the QC interface, the operator  can place the controls in the daily samples and complete the QC analysis in the  sample counting interface.

Before daily sample counting, when the operator edits the work order or enters  information in the "Next Sample" dialog box, the special "QC ID" that has been  set is entered as the "Sample ID".

According to the selected QC mode, select one of the following methods for QC  analysis:

• WB

• PD

1. Prepare controls in accordance with their MANUAL.

2. Sample preparation in WB mode and PD mode is carried out as described in   Section 5.5.1 Sample preparation.

3. When the counting operation is ready (i.e. the status icon and the indicator   light of the instrument are solid blue), the prepared sample is placed under the   sample probe and the QC analysis starts by clicking on the Aspirate Key.

4. After the aspiration, the operator can safely remove the control.

5. After the analysis, the QC results are automatically saved to its empty file, and

S

the newest QC results are displayed in the current interface.

M

A

R

G

6. If necessary, repeat the above steps to continue the QC analysis.

O

R

P

C

Q

7

E

H

R

T

E

T

G

P

N

A

I

H

S

C

U

4) After the aspiration, the operator can safely remove the control.

4

5

6

6

0

0

7

R

ET

PA

HC

Edit and save results (Administrator level)

Click the "Edit Result" button in the QC interface to edit the results. After finishing

the edit, press the “OK" button to save it. The edited result is automatically

marked with "E".

Restore results (Administrator level)

With Administrator privileges, the edited result can be restored to the initial

measurement value.

1. In the edit result interface, click the "Recover Result" button.

2. Select OK to restore the result.

3. Click "OK" to close the dialog and perform data recovery.

QC Results Review

After completing the QC analysis, the user can review the QC results in the

following two ways:

• QC Graph

• QC List

QC graph review

1. Click the “QC Graph” button in the “L-J QC Count” interface to enter the QC

graph interface corresponding to the QC file.

S

M

A

R

G

O

R

P

C

Q

E

H

T

G

N

I

S

U

2. Click the page up/page down buttons on the right of the QC graph to browse  the parameter QC results you wish to review. Click the page left/page right  buttons at the bottom of the QC graph to browse all the QC results.

QC list review

1. Click the "QC List" button in the "L-J QC Count" interface to enter the QC  graph interface shown below.

2. Click the page up/page down buttons on the right of the QC list to browse all  QC records. Click the page left/page right buttons at the bottom of the QC list to  browse all parameter results.

S

M

A

R

G

O

R

P

C

Q

7

E

H

R

T

E

T

G

P

N

A

I

H

S

C

U

6 60

7 60

7

R

ET

PA

HC

Delete (Administrator level)

1. Click the “Delete selected records" button to open the following dialog box.

2. Click "Yes" to delete the selected records.

• Delete operations are recorded in the log.

Print

To print the QC list, click the "Print" icon.

Export

To export the QC information and QC results of the current QC file, use the

following procedure:

1. Insert a USB flash drive and click the "Export" button.

2. The system will automatically detect the USB drive and export the data.

S

M

A

R

3. The system prompts the "Export Succeed" message.

G

O

R

P

C

Q

E

H

T

G

N

I

S

U

X-B QC

QC Principles

The X-B floating average method monitors the performance of the analyzer by  monitoring the stability of RBC parameters, such as MCV, MCH, and MCHC. It is  a QC method without controls. It, along with QC with controls, is the performance  monitoring method of the analyzer. They can reflect the analytical performance of  the analyzer from different aspects, and cannot replace each other.

The X-B method requires the use of random samples and therefore does not  apply to samples classified by disease. It involves a reference range consisting  of a given reference value and the upper and lower limits. The trend of the QC  results in the reference range is observed. This method is recommended when  the analyzer’s daily throughput is more than 100 samples.

The analyzer performs X-B QC on the three parameters of MCV, MCH,  and MCHC. The samples are the analyzer's normal count results, without  distinguishing between WB and PD modes. The number of samples for each X-B  numerical analysis set can be 20 - 200, and the analyzer can store up to 1000  X-B QC results. When the number of the QC result saved exceeds the limit, the  newest QC results will overwrite the oldest.

QC Settings

Click "Menu" > “QC" > "X-B QC" > "Setting" to enter the following X-B QC setting  interface.

S

M

A

R

G

O

R

P

C

Q

7

E

H

R

T

E

T

G

P

N

A

I

H

S

C

U

In the X-B QC Setting interface, you can edit the "X-B QC" information and the

8

9

6

6

“Ref. Value/Deviation setting” and perform "Sample validity setting".

0

0

7

R

ET

PA

HC

QC Analysis

After QC editing is completed, the system will automatically start performing X-B

QC counting.

Once 20 to 200 (according to the setting) valid sample results are obtained, the

system automatically executes an X-B QC calculation. The resulting QC results

can be reviewed in the X-B QC graph or the X-B QC list.

QC Results Review

After completing the QC analysis, the user can review the QC results in the

following two ways:

• QC Graph

• QC Lis

QC graph review

1. Click “Menu” > “QC” > “X-B QC” > “QC Graph” to enter the X-B QC graph

interface:

S

M

A

R

G

O

2. Select the file number of the QC file you want to review, and the interface

R

P

C

displays the file information and QC graph for the selected file.

Q

E

H

T

3. Click the page left/page right buttons at the bottom of the QC graph to browse

G

N

all the QC results.

I

S

U

QC list review

1. In the “X-B QC Graph” interface:

2. Click the "QC List" button to enter the QC list interface corresponding to the  QC file.

3.Click the page left/page right buttons at the bottom of the QC list to browse all  the QC records.

Similarly, the QC list also provides functions such as “Delete selected records"  and “Export".

S

M

A

R

G

O

R

P

C

Q

7

E

H

R

T

E

T

G

P

N

A

I

H

S

C

U

0 70

1 70

8

R

ET

PA

HC

Introduction

When to Calibrate

R

E

Z

Y

L

A

N

A

R

U

O

Y

G

N

I

T

A

R

B

I

L

A

C

CHAPTER 8

CALIBRATING

YOUR ANALYZER

The aim of calibration is to determine the deviation correction factor of blood  sample analysis under the specified conditions in order to obtain accurate  measurement results. To obtain accurate analysis results, the analyzer should be  calibrated according to the steps described in this chapter when necessary.

The analyzer provides three calibration methods: Manual calibration, calibration  with calibrators, calibration with fresh blood. The calibration modes include "WB"  and "PD".

All or part of the parameters of WBC, RBC, HGB, MCV, PLT, CRP can be  calibrated.

• Only operators with Administrator privileges can perform calibration.

• The operator shall use the calibrators and reagents specified by Zybio and store  and use them in strict accordance with their MANUAL.

• The calculation of reproducibility should also be included in the calibration step.

The analyzer has been calibrated before shipment. Since the analyzer itself is  stable in performance, there is no need for frequent calibrations. The operator still  needs to calibrate the analyzer in the following four situations:

• Before the first installation (usually by manufacture or an authorized  representative);

• After replacing main components;

• When there is obvious deviation in the QC data or the data exceeds the  predefined limit;

• When the main unit is not in use for a long time period and is put to use again;

How to Calibrate

• The analyzer must be calibrated, or the measured data cannot be used as valid  data.

Preparing Your Analyzer

Before calibration, check the analyzer according to the following steps to verify  that the background range, reproducibility and carry-over rate of the analyzer  are within the ranges specified in the MANUAL. Otherwise, you must find the  reasons and judge whether calibration is needed after the problem is solved. If  the problem cannot be solved, please contact the after-sales service department.

1. Check the main unit and reagents to ensure that the reagents are sufficient to   complete the entire calibration process. If the reagents are used up during the   calibration process, the calibration needs to be carried out again.

2. Perform background tests: Ensure that the background test results meet the   specified requirements (see Appendix A “Specifications” for the background   range).

3. Perform reproducibility tests: In the “Sample count" interface, count 10   consecutive times with a normal control or a blood sample equivalent to the   normal control range. In the “Review" interface, check the reproducibility of the   10 count results to ensure that they are within the specified range (see   Appendix A “Specifications” for the reproducibility indexes).

4. Detection of carry-over rate: Count 3 times with high value samples/controls,   and then immediately count 3 times with the compatible diluent/low value   samples. Then, the carry-over rate is calculated according to the following   formula.

Carry-over Rate= (j1-j3)/(i3-j3) ×100%

R

E

Z

Y

L

A

It is suggested that the operator establish a record file and make a record form

N

A

R

for archiving. The record form should include: date, source of calibrators, lot,  U

O

reference value and background value.

Y

G

8

N

I

T

R

A

E

R

T

B

P

I

A

L

H

A

C

C

2

3

• When the operating environment (such as temperature) has changed substantially.

7

7

0

0

8

R

ET

PA

HC

Manual Calibration

After the operator logs into the system with Administrator privileges, he/she can

click on the calibration factor of each parameter under the “Manual Cal.” interface

and enter and edit the new calibration factor.

Click “Management” > “Calibration” > “Manual Cal.” to enter the “Manual Cal.”

main interface as shown below. The calibration factor corresponding to each

parameter in the "WB" or "PD" mode and the operation time of the factor

are displayed in the interface. The operator selects and displays the current

calibration factor corresponding to the mode selected for manual calibration.

• The operator who logs in as a normal user can only view the calibration factor

in the current interface and cannot perform calibration. If he/she need to calibrate

the analyzer, he/she should first log out of the current user and log in as an

Administrator.

R

E

Z

Y

L

Use the following procedure to complete manual calibration.

A

N

A

The operator enters the “Manual Cal.” interface to view the calibration factor

R

U

and uses the following formula to calculate the new calibration factor for each

O

Y

parameter:

G

N

I

T

A

R

B

New Cal.Factor= Current Cal.Factor×Reference value

I

L

A

Mean measured value

C

If the calculated calibration factor of a parameter falls out of the effective range of  the calibration factor (the calibration range is 75% - 125%), then the calibration  factor is invalid. In this case, the operator must find the reason, troubleshoot,  recalibrate it, and calculate the calibration factor again. If the problem cannot be  solved, please contact the after-sales service or the authorized agent of Zybio .

After obtaining the new calibration factor, enter it in the calibration factor cell  where the parameter calibration is needed.

When the new calibration factor is entered, click “Save”.

Calibration with Calibrator

Click “Management” > “Calibration” > “Calibrator Cal.” to enter the interface  shown below.

R

• Calibration of calibrators can only be performed in the WB mode.

E

Z

Y

L

• The lot, expiry date and parameter reference value of the calibrator are shown

A

N

A

in the MANUAL of the calibrator.

R

U

O

Y

• The operator must use the calibrators designated by Zybio for this analyzer.

G

8

N

I

Zybio will not be responsible for any errory results caused by the use of other

T

R

A

E

R

T

calibrators.

B

P

I

A

L

H

A

C

C

4 70

5 70

8

R

ET

PA

HC

Complete the calibration of the calibrator as follows:

1. Verify the mode on the instrument control panel.

2. Enter the lot of the current calibrator in the “Lot“ edit box.

3. Set the expiry date. The default expiry date of the calibrator in the analyzer is

the current date. If you need to modify it, click the "Exp. Date" edit box to

set the expiry date. The expiry date of the calibrator cannot be earlier than the

current system date.

4. Enter the “Exp.date”. The entered expiration date should be either the

expiration date printed on the labeling or the open-container expiration date,

whichever is earlier. The open-container expiration date is calculated as

follows: the date that container is opened + the open-container stability days.

5. Enter the target value in the “Target" edit box corresponding to the parameter

to be calibrated.

6. Prepare the calibrator according to its MANUAL.

7. Press the Aspirate Key on the analyzer to start the calibration count.

8. When the total times of calibration counts reach n (n is greater than or equal

to 5), the analyzer will calculate the mean value, CV% and the new calibration

factor.

9. Save the calibration factor.

If the calculated calibration factor of any parameter to be calibrated is not

within the range of 75% - 125% (i.e. < 75% or > 125%), or the CV% value of

any calibration parameter exceeds the reproducibility index of the analyzer, the

calibration factor value will not be saved.

R

E

Z

Y

L

A

N

A

R

U

O

Y

G

N

I

T

A

R

B

I

L

A

C

Calibration with Fresh Blood

Click “Management”, “Calibration” and “Fresh Blood Calibration” to enter the  main “Fresh Blood Calibration” interface shown below.

Perform fresh blood calibration as follows:

1. Prepare 3 - 5 normal fresh blood samples according to the sample preparation   method introduced in Chapter 5 Routine operation.

2. Take the 3 - 5 samples of the prepared normal fresh blood, measure at least   5 times on a reference instrument, and calculate the mean value, which is used   as the reference value. Or, measure and calculate according to the reference   method, and the obtained data is used as the reference value.

3. Click the “Mode" button to select the fresh blood calibration mode and then   the WB or PD mode.

R

4. Select the number of the current calibration blood sample in the “Current

E

Z

Y

Blood Sample ID" drop-down list.

L

A

N

A

5. Select the parameter to calibrate from the check boxes in the first row of the

R

U

O

list.

Y

G

8

N

I

6. Enter the reference value of the parameter to be calibrated in the edit box

T

R

A

E

R

T

corresponding to “Reference value”.

B

P

I

A

L

H

A

C

C

7. Prepare WB or PD fresh blood samples.

6

7

7

7

0

0

8

R

ET

PA

HC

8. Place the blood sample under the sample probe and press the “Aspirate Key”

on the instrument to start the calibration counting sequence.

9. Once the calibration count is completed, the calibration count progress bar

dialog box closes automatically, and the analyzer will perform different

processing depending on the calibration count results.

• If the calibration count result is not within the linear range, but within the display

range, the calibration count result is displayed in the list and the calibration result

is not saved.

• If the calibration count result is not within the display range, the calibration

count result in the list will show * * * (* * * displayed according to the data format

of each parameter), and the calibration result will not be saved.

• If the calibration count result is in the linear range, it is valid and is displayed.

After obtaining valid calibration and count results, the check box in front of

them changes to "√", and they are used in the calculation of the blood sample

calibration factor by default.

10. For each blood sample, when 5 or more successive valid count results are

available, the CV% and calibration factor are calculated for each parameter.

11. Press the "Blood Sample 2" to "Blood Sample 5" buttons to enter the

"Fresh Blood Calibration" interface for Blood Samples 2-5. Follow the calibration

procedure for sample 1 and complete the calibration counts for at least three

more fresh blood samples to get their respective calibration factors.

12. After obtaining the calibration factors of more than 3 fresh blood samples,

press the "Calculate" button to enter the fresh blood calibration result

“Calculation” interface shown in the following figure.

R

E

Z

Y

L

A

N

A

R

U

O

Y

G

N

I

T

A

R

B

I

L

A

C

• Click the check box in front of each blood sample’s calibration factor to  select or cancel the calibration factors to be used in the calculation of the mean  calibration factor.

• When the “√” ticked calibration factors are no less than 3 sets, the CV% value  of the calibration factors will be automatically recalculated accordingly.

13. If you have not calculated the mean calibration factor, switch to the fresh   blood calibration interface, or while switching the calibration mode, there will   be a reminder” if the mean calibration factor has not been calculated, exit   and abandon all intermediate data. Continue or no?”

14. If the calculated mean calibration factor is within the valid range, the fresh   blood calibration interface is switched on.

CRP Calibration

On the menu bar, select "Management" > "Calibration" > "CRP Cal." to enter the  interface shown below.

R

E

Z

Y

L

A

N

A

Follow these steps to complete CRP calibration:

R

U

O

1. Select "Method" as Logit-Log4p.

Y

G

8

N

I

2. Enter the target of the calibrator in the target value field.

T

R

A

E

R

T

B

P

I

A

L

3. Three tests are performed using the calibrators of corresponding solubility and

H

A

C

C

the analyzer calculates the SD, CV, and mean values of the calibrator response.

8

9

7

7

0

0

8

R

ET

PA

HC

• The tests of the calibrator are from the low to the high concentrations, and the

order cannot be reversed.

• If the CV value of the calibration exceeds the set value of the system, the user

needs to recalibrate.

4. After testing the three concentrations of the calibrator, click "Calculate". The

analyzer will determine the correspondence between the CRP concentration

and the response according to the target and the calculated results of the

response.

5. (Optional) Click "Graphics View” to view the calibration curve; click "Close" to

close the window.

6. Click “Save” to save the current calibration result; or click “Empty” and click

“Yes” in the pop-up dialog box to clear the current calibration result and re-run

the calibration.

R

R

E

E

Z

Z

Y

Y

L

L

A

A

N

N

A

A

R

R

U

U

O

O

Y

Y

G

G

8

N

N

I

I

T

T

R

A

A

E

R

R

T

B

B

P

I

I

A

L

L

A

H

A

C

C

C

0 80

1 80

9

R

ET

PA

HC

Introduction

E

R

A

W

T

F

O

S

R

E

Z

Y

L

A

N

A

E

H

T

G

N

I

Z

I

M

O

T

S

U

C

CHAPTER 9

CUSTOMIZING THE  ANALYZER SOFTWARE

Initialization setting of the analyzer has been conducted before shipment. The  interface the user sees after the initial power-on is the system default. To meet  the different needs of actual applications, you can use the “Setup” program to  customize the software options as introduced in this chapter.

Click "Management" > "Settings" to enter the setting interface, as shown below:

General Settings

Auxiliary Settings

In the menu, select "Management" > "Settings" > "General" > "Auxiliary" to enter  the interface shown below.

• Next Sample Setup

Select sample ID entry method

Click the drop-down list and select the sample ID entry method from the following  options:

• Auto Increment

• Manual Entry

E

R

A

W

T

F

Locked prefixion digits

O

S

The user can set the number of digits in Sample ID that do not adopt auto

R

E

Z

increment.

Y

L

A

N

A

This edit box is activated when the Sample ID entry mode is "Auto Increment".

E

H

T

In the "Locked prefixion digits" edit box, enter the desired number n. The first n

G

N

9

I

Z

characters of all Sample IDs do not adopt the increment.

I

R

M

E

T

O

P

T

A

S

H

U

C

C

2 80

3 80

9

R

ET

PA

HC

• First sample ID setup after startup

The operator may customize the first sample ID after startup by entering it into

the edit box. Or the operator can select "Continue with Sample ID before last

shutdown".

• Warning Flags

Set up warning flags: The operator can select the suspect warning flags in the

drop-down list. The default is “?”.

Set high and low warning flags: The operator can enter single characters in the

two edit boxes or select high and low warning flags in the drop-down list (the

default high warning character is “↑” and the default low warning character is “↓”).

Print Settings

In the menu, select "Management" > "Settings" > "General" > "Print" to enter the

interface shown below.

E

R

A

W

T

F

O

S

R

E

Z

Y

L

A

N

Print setting steps are as follow:

A

E

1. Select the print device in the "Print device" drop-down box. There are two

H

T

G

types of print devices: “Recorder" and "Printer".

N

I

Z

I

M

2. Set paper type

O

T

S

Select the paper type in the "Paper type" drop-down box. The available paper

U

C

3. Set print title

Enter the report title in the "Report Title" box.

4. Set a report template

Select the type of print template in the "Template" drop-down box.

5. Set number of copies

In the “Number of copies” box, enter the number of copies to be printed for   each report.

6. Automatic print settings

The operator can choose automatic print as needed.

System Time Setting

From the menu, select "Management" > "Settings" > "General" > "Time" to enter  the interface shown below. The date, time and date format of the analyzer can be  set in this interface.

E

R

A

W

T

F

O

S

R

E

Z

Y

L

A

N

A

E

H

T

G

N

9

I

Z

I

R

M

E

T

O

P

T

A

S

H

U

C

C

sizes are A4, A5, Z4, continuous paper, letter paper, and thermal paper

4

5

8

8

0

0

9

R

ET

PA

HC

System Settings

E

R

A

W

T

F

O

S

R

E

Z

Y

L

A

N

A

E

H

T

G

N

I

Z

I

M

O

T

S

U

C

Laboratory Information Settings

Select "Management” > “Settings” > "General" > "Lab Information Settings" in the  menu. Enter the interface shown below. The operator can enter, save and view  laboratory information. The operator can click on the corresponding edit box and  enter relevant laboratory information as needed.

Automatic Maintenance Settings

Select "Management" > "Settings" > "System" > "Maintenance" in the menu.  Enter the interface shown below.

• Auto Sleep

If you need to set the time required to start auto sleep after the operations of   the fluidic components stop, it can be entered in the “Wait” edit box. The range   is 30 - 60 minutes.

• Auto Cleanser Soak

Select the start time of the probe cleanser maintenance. If you need to set the   probe cleanser maintenance time, enter it in the interval reminding time edit box.

Gain Settings ( for administrators only)

Select "Management" > "Settings" > "System" > "Gain" in the menu. Enter the  interface shown below.

E

R

A

• HGB gain

W

T

F

Adjust the HGB Blank to 4.5V ± 0.1V.

O

S

R

E

Z

Y

L

A

N

A

E

H

T

G

N

9

I

Z

I

R

M

E

T

O

P

T

A

S

H

U

C

C

6 80

7 80

9

R

ET

PA

HC

Parameter Settings Parameter Unit Settings

Select "Management" > "Setup" > "Parameter" > "Para. Unit" in the menu.

Enter the following interface.

• Select Unit

Click the "Select Unit" drop-down list and select the required unit.

• Customized unit settings

Under each unit system, the operator can click on the "Unit" cell to

customize the unit for any parameter. Click the "Restore Default" button to

restore the default settings for each unit.

Reference Range Settings

Select "Management" > "Setup" > "Parameter" > "Ref. Range" in the menu.

Enter the interface shown below.

E

R

A

W

T

F

O

S

R

E

Z

Y

L

A

N

A

E

H

T

G

N

I

Z

I

M

O

T

S

U

C

User Settings

The interface provides 5 internal reference groups and 10 customized reference  groups for the operator to select and set up. Each laboratory shall select  appropriate reference ranges according to their actual samples and set up  appropriate reference intervals. The reference interval varies according to race,  gender, age, and geographical location.

• Customized Group

In the reference group list, select the target reference group row and click the   “New" button to enter the reference group setting interface and set information   such as the name, age range, and parameter range of the reference group.

Select "Management" > "Settings" > "User" in the menu. Enter the interface  shown below.

E

R

A

W

T

• Reset PWD

F

O

S

The currently logged in user can modify the password of the current login

R

E

Z

account. Enter up to 12 characters.

Y

L

A

N

A

E

H

T

G

N

9

I

Z

I

R

M

E

T

O

P

T

A

S

H

U

C

C

8 80

9 80

9

R

ET

PA

HC

• New

1) Click the “New” button to bring up the following dialog box.

2) Enter relevant information such as “Username”, “Name” and “Password” in

each edit box. The user needs to enter the “Username” to log in. The name

is that of the "Inspector" and “Validator" as seen in the Review and printed

report.

3) Select user authority.

4) Click “OK” to save and close the dialog box.

• The username cannot be blank. A maximum of 12 characters are allowed.

E

E

R

R

• The password cannot be blank. A maximum of 12 characters are allowed.

A

A

W

W

T

T

F

F

• The name cannot be blank. A maximum of 20 characters are allowed.

O

O

S

S

R

R

E

E

Z

Z

Y

Y

L

L

A

A

N

N

A

A

E

E

• Delete a user

H

H

T

T

Click on the list to select a user and click on the “Delete" button to delete it.

G

G

N

N

9

I

I

Z

Z

I

I

R

M

M

E

T

O

O

T

P

T

S

A

S

U

H

U

C

C

C

0 90

1 90

0 1

R

ET

PA

HC

Introduction

R

E

Z

Y

L

A

N

A

R

U

O

Y

G

N

I

C

I

V

R

E

S

CHAPTER 10

SERVICING

YOUR ANALYZER

To ensure the accurate and effective performance of the analyzer, the operator  shall carry out routine maintenance according to the requirements of this chapter.  The analyzer provides multiple maintenance functions to help the operator to  complete the maintenance.

This chapter introduces various maintenance functions of the analyzer as well as  some measures in case of errors and alarms.

The surface of all parts of the analyzer pose potential biological hazards.  Therefore, safety precautions should be taken during operations and  maintenance.

• Improper maintenance may damage the analyzer. The operator must carry out  maintenance according to the MANUAL.

• If there is any issue not clearly mentioned in the MANUAL, please contact the  after-sales service department of Zybio and the professional personnel des ignated by Zybio for maintenance suggestions.

• The analyzer must be maintained with the spare parts provided by Zybio. If you  have any questions, please contact the after-sales service department of Zybio.

• When carrying out maintenance, avoid touching the sharp tip of the sample  probe.

The following is a list of tools that may be required in maintenance. SERIAL# TOOL

1. Phillips screwdriver

2. Slotted screwdriver

3. Allen wrench

Maintenance

Maintenance includes: Routine maintenance, cleaning and whole machine  maintenance.

On the menu bar, choose "Management" > "Service" > "Maintenance” and enter  the interface shown below.

• Unclogging

Unclogging includes burning and back flushing. In the event of clogging, the   unclogging operation can be performed. The operations are as follow:  1) Click the "Unclog” button to start unclogging.

2) After the unclogging is completed, the system prompts “Maintenance   completed”.

3) If necessary, a single channel operation of “WBC Flush Unclog” and “RBC   Flush Unclog” can be performed.

• Replace Reagent

The reagents for the corresponding channels can be refilled.

R

E

Z

Y

L

• Empty

A

N

A

Empty the liquid from the corresponding channel.

R

U

O

0

Y

1

G

R

N

I

E

C

T

I

P

V

A

R

H

E

C

S

4. Medical gloves

2

3

9

9

0

0

5. Alcohol

0 1

R

ET

PA

HC

Self-check

R

E

Z

Y

L

A

N

A

R

U

O

Y

G

N

I

C

I

V

R

E

S

• Cleaning

The user needs to clean the parts in the following situations:  1) If the WBC and/or HGB background results exceed the background   ranges, the WBC chamber can be cleaned.

2) If RBC and/or PLT results exceed the background ranges, the RBC   chamber can be cleaned.

3) If the sample probe is dirty, perform sample probe cleaning.

• Probe Cleanser Maintenance

The user should perform Probe Cleanser Deeply Soak under the following   conditions:

1) If the background is out of range and the QC result is abnormal because   the analyzer has not been in use for a long period, or if unclogging fails   despite other maintenance operations.

2) If the instrument shuts down due to abnormal power failure.

• Whole device

1) Whole device initialization: Restore all moving parts and sensors of the   instrument to their initial state.

2) Whole device cleaning: Clean all the fluidic components of the instrument.  3) Prime: Fill the instrument’s fluidic components with reagent.  4) Drain All: When the instrument has not been used for more than one   week, perform “Drain All” by emptying the instrument and washing the   instrument with distilled water according to the interface prompts.

Select "Management" > "Service" > "Self-check" to enter the following interface  to perform system and valve self-checks.

System

Calibration Log

Choose "Management" > "Service" > "System Cal." to enter the following  interface to perform touch screen calibration.

Choose "Management" > "Service" > "Log" to enter the following interface.

R

E

Z

Y

L

A

N

A

R

U

O

0

Y

1

G

R

N

I

E

C

T

I

P

V

The error, parameter modification, and daily operation logs can be viewed in this

A

R

H

E

C

S

interface.

4 90

5 90

0 1

R

ET

PA

HC

Status

R

E

Z

Y

L

A

N

A

R

U

O

Y

G

N

I

C

I

V

R

E

S

The logs can be used to record the use of the analyzer, and are important  for the operator to search the use history and for the service personnel to  troubleshoot problems.

• Log export

1) Insert a "USB flash drive".

2) Click "Export" and the following dialog box will pop up.  3) Select the log records to export.

4) The interface prompts “Export succeeded”.

Version

Select "Management" > "Status" > "Version" to view the instrument’s software  version.

Sensor

Select "Management" > "Status" > "Sensor" to view the instrument’s sensor  status.

Voltage

Select "Management" > "Status" > "Voltage" to view the instrument’s voltage  status.

R

E

Z

Y

L

A

N

A

R

U

O

0

Y

1

G

R

N

I

E

C

T

I

P

V

A

R

H

E

C

S

6 90

7 90

CHAPTER 11

TROUBLESHOOTING

ERROR NAME ACTIONS

1. Click “Clear”, and this error will automatically clear.

Communication Abn

2. If the error still exists, please contact our after-sales   service department.

Introduction

YOUR ANALYZER

This chapter describes the possible errors of the analyzer and provides the  corresponding corrective actions.

Voltage Abnormal 1. Please turn off the power directly and contact our after-  sales service department.

System Clock Abnormal 1. Please turn off the power directly and contact our after-  sales service department.

1. Click “Clear" and enter the barcode of the new reagent in

the popup message.

Error Information

and Handling

R

E

Z

Y

L

A

N

A

R

U

O

Y

G

N

I

T

Samples, controls, calibrators, waste liquid, etc. pose potential biological  hazards. When the operator comes into contact with related articles in the  laboratory, he/she shall comply with the laboratory safety practices and wear  personal protective equipment (such as laboratory protective clothing, gloves,  etc.).

This MANUAL does not serve as the maintenance manual and it only provides the  measures that the operator should take when the analyzer gives a error alarm.

During the use of the analyzer, if an abnormal condition is detected, the  corresponding error prompt message will be displayed at the bottom of the  analyzer’s display interface, and the main unit will also sound an alarm.

Click on the error alarm area to open the error dialog box. The error dialog box  provides the error messages and help information. The error messages will be  displayed in the chronological order in which the errors occur.

The operator can select the error message in the dialog box by clicking on it. The  help information of the selected error can be viewed in the “Error help" list box at  the bottom of the dialog box. The help information of the first error is displayed  by default. The operator shall deal with the errors in sequence according to the

No Reagent R2 Diluent Expired

Lyse Expired Waste Full

Diluent Empty Lyse Empty

2. Change the reagent bottle and click the "Apply" button to   prime the reagent.

3. If the error still exists, please contact our after-sales service   department.

1. Click “Clear" and enter the barcode of the new reagent in   the popup message.

2. Change the reagent bottle and click the "Apply" button to   prime the reagent.

3. If the error still exists, please contact our after-sales

service department.

1. Click “Clear" and enter the barcode of the new reagent in   the popup message.

2. Change the reagent bottle and click the "Apply" button to   prime the reagent.

3. If the error still exists, please contact our after-sales

service department.

1. Empty the waste bucket or change to a new one.

2. Click the “Clear" button and this error will automatically   clear.

3. If the error still exists, please contact our after-sales

service department.

1. Click “Clear" and enter the barcode of the new reagent in   the popup message.

2. Change the reagent bottle and click the "Apply" button to   prime the reagent.

3. If the error still exists, please contact our after-sales

service department.

R

E

Z

1. Click “Clear" and enter the barcode of the new reagent in

Y

L

the popup message.

A

N

2. Change the reagent bottle and click the "Apply" button to

A

prime the reagent.

R

U

3. If the error still exists, please contact our after-sales

O

Y

service department.

G

N

I

T

1 1

R

ET

O

O

contents of the error help.

H

S

E

L

To help the operator look up errors, error messages that the analyzer may display

B

Voltage Abnormal 1. Please turn off the power directly and contact our after-

O

sales service department.

O

1

H

1

S

E

R

L

1. Click “Clear” and this error will automatically clear.

E

B

T

PA

HC

U

O

are listed in the MANUAL, in which the possible causes and corrective actions

R

T

Syringe Component Abnormal

2. If the error still exists, please contact our after-sales

P

U

A

O

service department.

H

R

C

T

are also provided. Thus, the operator is able to troubleshoot and clear the error

8

9

9

9

messages accordingly. If the problem still exists, please contact the after-sales

0

0

service department.

ERROR NAME ACTIONS

1. Click “Clear” and this error will automatically clear.

Repair

The instrument may fail during use. If the user cannot repair it by himself/herself,  he/she need to contact a service technician to check the failure on site. It may be  necessary to replace the accessories. See Appendix C for a list of accessories.

Sampling Component Abnormal Background Abnormal

HGB Blank Voltage Abnormal

Vacuum Pressure Abnormal WBC Clog

WBC Aperture Voltage Abnormal WBC Impedance Signal Interference RBC Clog

RBC Aperture Voltage Abnormal

2. If the error still exists, please contact our after-sales   service department.

1. Click “Clear” and this error will automatically clear. 2. If the error still exists, please contact our after-sales   service department.

1. Please check id the diluent empty or not. If there is no   reagent, please replace it with a new one. 2. Click the “Clear" button and this error will automatically   clear.

3. If the error still exists, please contact our after-sales   service department.

1. Click “Clear” and this error will automatically clear. 2. If the error still exists, please contact our after-sales   service department.

1. Click “Clear” and this error will automatically clear. 2. If the error still exists, please contact our after-sales   service department.

1. Click “Clear” and this error will automatically clear. 2. If the error still exists, please contact our after-sales   service department.

1. Please remove the source of interference. 2. Click the “Clear” button and this error will automatically   clear.

1. Click “Clear” and this error will automatically clear. 2. If the error still exists, please contact our after-sales   service department.

1. Click “Clear” and this error will automatically clear. 2. If the error still exists, please contact our after-sales   service department.

1 1

R

ET

PA

HC

RBC Impedance Signal Interference 1. Please remove the source of interference.

2. Click “Clear” and this error will automatically clear.

R

R

1. Click “Clear” and this error will automatically clear.

E

E

Z

Z

2. If the error still exists, please restart the analyzer.

Y

Y

L

L

Detection Abnormal

A

A

3. If the error still exists, please contact the after-sales

N

N

service department.

A

A

R

R

U

U

O

O

Y

Y

G

G

N

N

I

I

T

T

O

O

O

O

1

H

H

1

S

S

E

E

R

L

L

E

B

B

T

P

U

U

A

O

O

H

R

R

T

C

T

0 01

1 01

Classification

Reagents

Applicable Tubes

APPENDIX A

SPECIFICATIONS

According to the CE classification, Z3/Z3 CRP Series belongs to In vitro  diagnostic medical devices other than those covered by Annex II and devices for  performance evaluation.

Please use the reagents produced by Zybio.

According to the CE classification, Z3/Z3 CRP Series belongs to In vitro  diagnostic medical devices other than those covered by Annex II and devices for  performance evaluation.

TYPE SPECIFICATIONS AND DIMENSIONS APPLICABLE MODE Vacuum blood collection tube ∅12-15×75mm(without tube cap) WB mode ∅10.7×42mm(size without cap), 0.5ml, it

Parameters

PARAMETER ABBREVIATION DEFAULT UNIT WBC count WBC 109/L Lymphocyte count Lym# 109/L Intermediate cell count Mid# 109/L Neutrophil count Gran# 109/L Percentage of lymphocytes Lym% % Percentage of intermediate cells Mid% % Percentage of neutrophils Gran% % RBC count RBC 1012/L Hemoglobin HGB g/L Mean corpuscular volume MCV fL Mean corpuscular hemoglobin content MCH pg

Mean corpuscular hemoglobin

concentration MCHC g/L RBC distribution width - coefficient of

variation (RDW-CV) RDW-CV % RBC distribution width - standard

deviation (RDW-SD) RDW-SD fL Hematocrit HCT % Platelet count PLT 109/L Mean platelet volume MPV fL Platelet distribution width PDW / Plateletcrit PCT %

Small anticoagulant tube

can be tested for cap opening.

Recommended: 0.5ml closed anticoagulant  tube (REF. 365974) produced by BD Inc.

Capillary WB mode

Platelet-large cell ratio P-LCR % Platelet-large cell count P-LCC 109/L

A

XI

D

NE

PPA

Centrmanualge tube (bullet) ∅11×40mm 0.5ml and 1.5ml centrmanual

ge tubes

S

N

O

I

T

A

C

I

F

I

C

E

P

S

PD and capillary WB  modes

C-reactive protein CRP mg/L

Hypersensitive C-reactive protein Hs-CRP mg/L

Sampling Features Sample Volume Required for Each Analysis

WB mode ≤ 15μl

PD mode ≤ 20μl

S

N

O

I

A

T

A

X

I

C

I

D

F

I

N

E

C

P

E

P

P

A

S

2 01

3 01

Throughput Throughput

CBC MODE CBC MODE

WB mode Throughput no less than 70 samples/hour  WB mode Throughput no less than 70 samples/hour  PD mode Throughput no less than 70 samples/hour

• Reproducibility

PARAMETER DETECTION RANGE WB (CV) PD (CV) (7.00 - 15.0)×109/L ≤ 2.0%

WITH CRP

PD mode Throughput no less than 70 samples/hour  WB mode Throughput no less than 60 samples/hour  WB mode Throughput no less than 60 samples/hour

WBC

≤ 4.0%

(4.00 - 6.9)×109/L ≤ 2.5%

WITH CRP  MODE

MODE

PD mode Throughput no less than 60 samples/hour  PD mode Throughput no less than 60 samples/hour

RBC (3.50 - 6.50)×1012/L ≤ 1.5% ≤ 3.0% HGB 110 g/L - 180 g/L ≤ 1.5% ≤ 3.0% MCV 70fL - 110fL ≤ 0.5% ≤ 2.0%

Performance Specifications

• Background/Blank Test

PARAMETER BACKGROUND /BLANK COUNT REQUIREMENTS WBC ≤0.2 × 109/L

RBC ≤0.02 × 1012/L

PLT

CRP

• Carryover

(100 - 149)×109/L ≤ 5.0% ≤ 10.0% (150 - 500)×109/L ≤ 4.0% ≤ 8.0% (0 - 10) mg/L ≤ 10% ≤ 15% (10.1 - 320) mg/L ≤ 5.0% ≤ 10%

HGB ≤1g/L

HCT ≤0.5%

PLT ≤5 × 109/L

CRP ≤0.2mg/L

• Linearity Ranges

PARAMETER LINEARITY RANGE LINEAR ERROR (WB MODE)

(0 - 100.00)×109/L ≤±0.30×109/L or ≤±5%

WBC

(100.01 ~ 300.00)×109/L ≤±9%

RBC (0 - 8.00)×1012/L ≤±0.05×1012/L or ≤±5% HGB (0 - 250)g/L ≤±2g/L or ≤±2%

(0 - 1000)×109/L ≤±10×109/L or ≤±10%

PLT

(1001 - 4000)×109/L ≤±12%

HCT 0% - 67% ≤± 2% (HCT)  or ≤± 3% (error percentage)

0.2 mg/L - 10.0 mg/L ≤±1.0mg/L

PARAMETER CARRYOVER WBC ≤ 0.5%

RBC ≤ 0.5%

HGB ≤ 0.5%

PLT ≤ 1.0%

CRP ≤ 0.5%

• Comparability

ITEM COMPARABILITY DEVIATION WBC ≤± 5%

RBC ≤± 2.5%

HGB ≤± 2.5%

PLT ≤± 8%

MCV ≤± 3%

A

XI

D

NE

PPA

CRP

S

N

O

I

T

A

C

I

F

I

C

E

P

S

10.1 mg/L - 100.0mg/L ≤±15% 100.1 mg/L - 320.0 mg/L ≤±20%

Input and

Output Devices

Touch Screen

10.4-inch TFT color touch screen, up to 24-bit color, resolution: 800 ×600.

S

Indicator Light

N

O

I

A

Used to indicate the analyzer's status: Power On/Off, Running, or Sleep.

T

A

X

I

C

I

D

F

I

N

Recorder

E

C

P

E

P

P

Built-in thermal recorder.

A

S

4

5

0

0

Buzzer

1

1

Used to indicate that the instrument is malfunctioning. The buzzer’s alarm sound

can be cleared automatically by tapping the touch screen or “Clear”.

Main Unit Interfaces Sound Pressure

Main Unit Interfaces

• One network port, one built-in network card with networking function,

compatible with TCP/IP protocol

Sound Pressure

A

XI

D

NE

PPA

Power Supply

Fuse

Electromagnetic

Compatibility

S

N

O

I

T

A

C

I

F

I

C

E

P

S

• Four USB ports

VOLTAGE INPUT POWER FREQUENCY

MAIN UNIT 100V - 240V ≤200VA 50Hz

Analyzer fuse specifications: T 6.3AH 250V

• Always use a fuse meeting the specified specifications.

The fuse can be replaced by the user. To replace the fuse, disconnect the power  cord and pull the fuse out of the fuse slot in the filter:

Fuse

Fuse slot

1. Do not use this product near strong radiation sources (such as unshielded RF   sources). Doing so may interfere with the normal operation of the product.

2. This product meets the emission and immunity requirements as specified in   GB/T 18268.1 and GB/T 18268.26.

3. This product is designed and tested in accordance with GB 4824 Class A   equipment.

Operating

Environment

Dimensions

and Weights

Contraindications Safety Classification

ENVIRONMENTAL REQUIREMENTS STORAGE OPERATION

Ambient temperature -10℃-40℃ 15℃-35℃

Relative humidity 10%-90% 20%-85%

Atmospheric pressure 50kPa-106kPa 70kPa-106kPa

• Be sure to store and use the analyzer under the specified environmental  conditions.

MAIN UNIT

Width (mm) ≤300

Height (mm) ≤410 (with support pads)

Depth (mm) ≤400

Weight (Kg) ≤20kg

None

Level of transient overvoltage: Category II

S

N

O

I

A

Rated pollution degree: Level 2

T

A

X

I

C

I

D

F

I

N

E

C

P

E

P

P

A

S

6

7

4. Before using this product, users need to evaluate the electromagnetic

0

0

1

1

environment.

APPENDIX B APPENDIX C KEY PARTS LIST OF SPARE PARTS

B

XI

D

NE

PPA

SERIAL# NAME

1 Switch power supply

2 Power cord

3 Filter

4 Fuse

5 Stepper motor

S

T

R

A

P

Y

E

K

SERIAL# PART

1 Switch power supply

2 Power cord

3 Filte

4 Fuse

5 Stepper motor

6 Three-way solenoid valve

7 Two-way solenoid valve

8 WBC counting chamber

9 RBC counting chamber

10 Liquid pump

11 Air pump

12 Main control board PCBA

13 Driver board PCBA

14 CRP measuring component

15 Open sample probe

16 Swab

17 Touch screen

18 Display

S

T

R

A

P

E

R

C

A

P

X

I

S

D

F

N

O

E

T

P

S

P

I

A

L

8 01

9 01

APPENDIX D

TOXIC AND HAZARDOUS  SUBSTANCES OR

ELEMENTS

TOXIC AND HAZARDOUS SUBSTANCES OR ELEMENTS

PART Built-in

LEAD (PB) MERCURY  (HG)

CADMIUM  (CD)

HEXAVALENT  CHROMIUM  (CR(VI))

POLYBRO MINATED  BIPHENYLS  (PBB)

POLYBRO MINATED  DIPHENYL  ETHER

(PBDE)

circuit board × 〇 〇 〇 〇 〇 Enclosure × 〇 〇 × 〇 〇 Display × 〇 〇 〇 〇 〇

S T

N

Optoelec tronic com ponents

Internal

electronic  wires

× 〇 〇 〇 〇 〇 × 〇 〇 〇 〇 〇

D

XI

D

NE

PPA

Accessories × 〇 〇 〇 〇 〇

E

M

E

L

E

R

〇: Indicates that the contents of the toxic and hazardous substances in all homogeneous materi

O

S

E

als of the part is below the limit specified in SJ/T 11363-2006.

C

N

×: Indicates that the contents of the toxic and hazardous substances in at least one homogeneous

A

T

S

B

material of the part exceeds the limit specified in SJ/T 11363-2006.

U

S

S

U

P/N: 70-01-0002-00[1.0]

O

D

R

A

Z

A

H

D

N

A

C

I

X

O

T

0 1

1